Learner Name:		Gareth Wyn Thomas

Course:			HNC Instrumentation Engineering Year 2

Module Code:			PCEE4009

Module Name:		Instrumentation, Automation and Control

Lecturer:			Jason Pointer

Contents

Task 1		Temperature Sensor Research						Page 3

Task 2		Testing the sensors							Page 10

Task 3		Temperature Circuit Design						Page 16

Task 4		Signal Transmission							Page 19

References										Page 20

Appendix	Tabulated data for Task 2						Page 21

Task 1 – Temperature Sensor Research

The two temperature sensors to be investigated are the Platinum RTD (Resistance Temperature Detector) and NTC Thermistor (Negative Temperature Coefficient). Both rely on the physical principle of the variation of resistance with temperature and are classified according to whether the material used in their construction is a metal or semiconductor.

- Platinum RTD.

Temperature sensor relies on the principle that resistance of a metal increases with temperature. This relationship can be represented by the non-linear equation: -

RT = R0 (1 + α1T + α2T2 +……       αnTn)

However, resistance of some metals increases in a reasonably linear way over a limited temperature range (e.g., platinum, copper, nickel and tungsten) as shown in Figure1.

Figure 1. Typical resistance – temperature characteristic of metals (Morris, 2006).

Therefore, the non-linear equation above may now be represented by the equation: -

RT = R0 (1 + αT)         where  RT is resistance at T0C

R0 is resistance at 00C

α is constant for the metal and called the

coefficient of resistance

Figure 1 indicates that platinum has the most linear resistance-temperature characteristic making it the preferred metal to use for a RTD for most applications. Research shows that its resistance-temperature relationship is linear within + 0.4% over a temperature range -2000C and +400C. Even at 1000C its accuracy is approx. + 1.2%.

The construction of a platinum RTD can take two forms. The first type called a thin-film RTD. It consists of a thin film of platinum paste deposited on a white ceramic substrate as shown in Figure 2.

Figure 2. A Resistive Temperature Detector (Morris, 2006).

The second type of design makes use of a platinum wire coil wound on a simple former (Figure 3).

Figure 3. Construction of a fine wire platinum resistance thermometer using fine (25µm diameter) platinum wire. Dimensions in mm (Harrison, 2015).

This former is housed in a protective case as shown in Figure 4.

Figure 4. Showing the former housed in protective case (Bird, May 1983).

Figure 4 illustrates a typical platinum RTD having a platinum wire coil wound on a simple former which can be made of metal, plastic, mica or ceramic housed in a metal (glass or porcelain) protective case or sheath. The platinum coil leads are connected to copper extension leads at points A and B. (Copper is used as platinum is a more expensive metal). Insulation spacers are included in the design to prevent the copper leads interfering with the housing. Dummy copper leads are sometimes introduced (especially if the copper extension leads B is too close to the heating source being measured) to eliminate errors. The copper extension leads (and dummy copper leads if included) are connected to an electrical circuit usually a Wheatstone Bridge or a potential divider circuit that will convert resistance to a voltage which is then calibrated to a temperature scale (Morris, 2006).

- NTC Thermistor.

This temperature sensor relies on the principle that as temperature is increased its resistance rapidly decreases in a non-linear manner. This is illustrated when comparing Figure 1 with Figure 5 which shows typical resistance – temperature characteristics of thermistors of different materials (nickel, copper and platinum).

Key:   Nickel ----     Copper - - - -  Platinum ……..

Figure 5. Typical resistance -temperature characteristics of thermistor materials (Morris, 2006).

As shown in Figure 5 each curve has a negative gradient. The gradient value between two temperature points represents the constant or coefficient value denoted by Beta (B) (as shown in the equation below) and depends on the metal from which the thermometer is made. Such thermistors are known as Negative Temperature Coefficient Thermistors (NTC). The value of B may range from approximately 3000 to 5000 and can be obtained from manufacturers data sheets. The nominated equation for such thermistors is shown below.

Figure 6. Thermistor Equation (Morris, 2006).

The construction of a NTC thermistor takes the form of a solid temperature sensor, consisting of a mixture of sensitive semi-conductor-based metal oxides such as chromium, cobalt, copper, manganese and nickel etc. These sensors can be formed into small ceramic shapes each connected to two leads. Figure 7 shows a rod, disc and bead shaped thermistor.

Figure 7. Showing a rod, disc and bead shaped thermistor (Bolton, 2015).

The principle of operation of a Platinum RTD and NTC Thermistor rely on the principle that their electrical resistance changes with temperature variation. Since resistance can be measured accurately it enables the sensor device to be used to measure temperature. As both devices are resistive, then according to Ohm’s Law, when a current passes through the device a voltage drop will occur which can be measured and then converted to a temperature (Morris, 2006).

One method of converting a resistance change into voltage change is use of a Wheatstone Bridge circuit as shown in figure 8. This signal conditioning is widely used as it can measure accurately a small change in temperature. It is often referred to as a Null Type Bridge and is used for calibration purposes (Morris, 2006).

Figure 8. Wheatstone Bridge Circuit (Morris, 2006).

With reference to figure 8

- The circuit comprises of 4 resistors and a galvanometer which is a sensitive zero micro ammeter.

- Resistor Rs (sensor) is the value to be measured and compared with other resistors of known value

- Resistors R2 and R3 have fixed and known values and Rv is a variable resistor

- A DC voltage V1 is applied across points A and C and the variable resistor RV is adjusted until the voltage measured across points B and D is zero (known as the null point).

- It can be shown using Ohm’s Law:-

V0 =   V1Rs      -   V1Rv

(Rs + R3)   (RV + R2)

V0 = V1  X       Rs      -         Rv

(Rs + R3)   (RV + R2)

At null point V0 = 0

Rs         =       Rv

(Rs + R3)     (RV + R2)

Inverted becomes

(Rs + R3)  =   (RV + R2)

Rs                  Rv

R3 = R2

Rs     Rv

Rs = R3 Rv

R2

But R2 = R3

So Rs = Rv

Thus, the value of the resistor Rs equals value of the known variable resistor.

So, if Rs represents a PRTD sensor its resistance at any temperature may be measured.

A deflection Type Bridge is sometimes used and differs from a Wheatstone Bridge in that the variable resistor Rv (Figure 8) is replaced by a fixed resistor R1 (Figure 9) of the same value as the normal value of the unknown resistor Rs.

As the resistance Rs changes then the output voltage V0 changes and the relationship between V0 and Rs must be calculated.

Figure 9. Analysis of Wheatstone Bridge (Morris, 2006).

The relationship between Vo and Rs is given by the equation

Vo  =   V1 x        Rs      -      R1

Rs + R3          R1 + R2

Figure 10. Wheatstone Bridge including dummy leads (Bird, May, 1983).

Figure 10 shows a Wheatstone Bridge including dummy leads. The dummy leads are included in the arm B and C to eliminate errors caused by the extension leads which are normally necessary in PRTD sensors. RƟ  represents the sensor.

In the case of thermistors, use is sometimes made of a potential divider circuit to convert the resistance change into voltage change then to a meter related to the temperature. Figure 11 illustrates a typical Potential Divider Circuit.

Figure 11. Potential divider circuit (Bird, May 1983).

Operational Amplification

When the signal output level of a measurement sensor is considered to be too low the signal needs to be amplified. This is achieved by an optical amplifier (op-amp). A simple op-amp is shown in Figure 12.

Figure 12. Simple op-amp circuit (Morris, 2006).

With reference to Figure 12.

- This is an electronic device that has two input terminals and one output terminal. The two input terminals are known as an inverting input (-) and non-inverting input (+).

- The input signal V1 is connected to the inverting input through the resistor R1 and the non-inverting input connects to ground.

- A feedback path is provided from output terminal through a resistor R2 to the inverting input terminal.

- The processed signal V0 at output is related to V1 (input) by the equation below.

V0 = R2 x Vin

R1

This signal amplification is defined by values of resistors R1 and R2 and the ratio R2/R1 is

known as amplification gain (or closed loop gain). The sign of the output signal V0 is set up as shown in Figure 12 is inverted if necessary. This can be corrected by feeding this signal through a further set up for unity gain (i.e., R1 = R2) this inverts the signal returning it to its original sign.

Task 2 – Testing the Sensors

To translate temperature into a measurable signal the most commonly used sensors are the Resistive Temperature Detector (RTD) and Thermistor. Each has its own operating principles, benefits, considerations, and drawbacks. In addition to the characteristics of the sensors themselves - operating range, sensitivity, linearity, response time - consideration must be given to the requirements each sensor type imposes on the measurement hardware.

The successive comparison of one instrument with another ultimately leads to the need for a datum instrument (Morris, 2006), which in this experiment is the IC temperature sensor which is that currently used by the company. The company required myself to investigate accuracy of the other two temperature sensors (RTD and Thermistor) against the IC datum instrument.

The values tabulated (see appendix) were used to plot the graph (Graph 1) of the RTD temperature sensor and the IC temperature sensor of Temperature (0C) against Time (secs) are calculated using the formulae below.

TRTD  =    VRTD – 100                 where  TRTD = temperature of RTD sensor in 0C

α                                      VRTD = voltage of RTD sensor in mV

α    = constant coefficient for platinum (0.385)

TIC = VIC x 100 -273                  where TIC  = temperature of IC sensor in 0C

VIC = voltage of IC sensor in volts

% error = TIC – TRTD  x 100

TIC

Example of calculations at time 5 seconds   VRTD = 108.11 mV   and α= 0.385

TRTD = VRTD - 100

α

= 108.1 -100

0.385

TRTD  = 21.03896 0C

TIC = VIC x 100 – 273

= 2.9863 x100 – 273

TIC  =25.63 0C

% error = TIC -TRD   X 100

TIC

= 25.63 -21.04   X 100

25.63

% error = 17.91%

The values tabulated (see appendix) were used to plot the graph (Graph 2) of the Thermistor temperature sensor and the IC temperature sensor of Temperature (0C) against Time (secs) are calculated using the formulae below.

R2 = ( 5314 x 5) - 5314

VT

T2 =      - 4050 /( ln (5314/R2) – (4050/298.15) )    - 273.15

TIC = VIC X 100 – 273

% error = TIC – T2  x 100

TIC

Example of calculations at time 5 seconds   VT = 2.499V

R2 =  ( 5314 x 5) - 5314

VT

=   5314 x 5     - 5314

2.499

R2  =   5318.25 Ω

T2 =  - 4050 /( ln (5314/R2) – (4050/298.15) )    - 273.15

= - 4050/ (ln(5314/5318.25) – (4050/298.15))   - 273.15

=  - 4050/ (-0.000799 – 13.5838)       - 273.15

= - 4050/(-13.58499)     -273.15

=  298.132 -273.15

T2 = 24.98 0C

TIC = VIC  x 100 –  273    where VIC = 2.997 V

= 2.997 x 100 – 273

= 299.7 - 273

TIC =  26.7 0C

% error = TIC – T2   X 100

TIC

=  26.7 – 24.982   X 100

26.7

= 0.06434 x 100

% error  = 6.434%

Data was gathered for the temperature recorded by a platinum Resistive Temperature Detector (RTD) as a low-level heating element over a period of 600 seconds in 1 second intervals using a Pico Scope and a PC as shown in the photographs below.

Photograph 1. Showing the experiment setup using a Pico Scope to record logged values on PC.

Photograph 2. Showing a close up of the NTC Thermistor and Platinum RTD.

A graph of the recorded temperature (0C) against time (seconds) was plotted and is shown by the blue coloured line on Graph 1. On the same sheet a graph of temperature, measured in 0C, by an IC temperature sensor (which is that currently used by the company) was also plotted against time (seconds) over the same period of 600 seconds in 1 second intervals. This data is represented by the red coloured line on Graph 1.

Graph 1.    IC temperature Sensor (Red), Platinum RTD Temperature Sensor (blue)

Temperature   (0C)

Time (seconds)

A second set of data was collected using the NTC Thermistor in 1 second intervals over a period of 600 seconds. The data gathered of temperature (0C) against time (seconds) is shown graphically as the blue coloured graph on Graph 2. The temperature (0C) against time in seconds recorded by the IC temperature sensor gathered in 1 second intervals over the same period of 600 secs is plotted as the red coloured graph on Graph 2.

Details of the individual readings are given in the Appendix of this Assignment.

Graph 2.   IC Temperature Sensor (red), NTC Thermistor Temperature Sensor (blue)

Temperature  (0C)

Time (seconds)

The response of an instrument to the parameter sensed, in this case temperature, can be characterized in several ways, how large the actual response is, how quickly the instrument responds, how reliable the readings obtained are and over what range of the input parameter the instrument functions reliably (Harrison, 2015).

The most fundamental use of any sensor or instrument as a detector is to determine the presence or absence of a change in a parameter and also the resolution which describes this smallest detectable change (Harrison, 2015). The quantitative response to the change is its sensitivity which is an absolute quantity and is the smallest absolute amount of change that can be detected by a measurement. It is the ratio of the response in the instrument to a unit change in the parameter sensed. i.e., the gradient of the graph.

Graph 1 shows that the two instruments follow broadly the same path and there is little difference between their gradients as time is increased during the 525 seconds. After this the IC graph appears to continue in an upward path whilst the RTD graph peaks after 375 seconds to follow a flatter path.

The graphs of the RTD and IC against time (Graph 1) have little difference between the readings at each interval of time. Errors in the range 19% to 10% over the first 30 seconds but becoming fairly consistent in the range 10% to 0.2% for the remaining 550 seconds.

The stability of a temperature sensor is an indication of its ability to maintain a consistent output at a given temperature. Material plays a key role in the stability of a given sensor. RTDs are often constructed of platinum for this reason as well as to ensure low reactivity (Harrison, 2015).

The graphs of data measured using the Thermistor and IC Temperature sensor are plotted on Graph 2. The red graph being that for the IC thermometer and the blue graph being that for the Thermistor. The two curves on Graph 2 show a fairly consistent gradient along the interval of time considered and this suggests consistent sensitivity. However there appears to be an increasing error between the readings from an initial 6.5% to a gradually regular increase in percentage error to almost 20% at the end of the time interval considered.

The instrument selected must be such that the range (i.e., the maximum and minimum values of a quantity that the instrument is designed to measure (Harrison, 2015)) is appropriate to the spread of values being measured in order that the best possible accuracy is maintained in instrument readings. Usually, the inaccuracy measurement rather than the accuracy figure is quoted to provide detail to the extent to which a reading might be wrong. This is quoted as a plus + or a minus - percentage of the full scale reading of the instrument (Morris, 2006). The error readings are noted in the previous paragraph.

A further important consideration is that of reliability of the measurement of the temperature reading and that the measuring equipment will operate to an agreed level of performance. That is that the system gives a particular accuracy for a specified period. The reliability of a measurement system is likely to change with time resulting from wear on contacts and general damage due to environmental conditions (Bolton, 2015).

Thus, it would be better for the company to adopt the RTD thermometer rather than the Thermistor to ensure a more consistent result. RTD thermometers tend to be fragile and can be easily damaged but the use of platinum, which is a very stable metal and chemically inert (even though its sensitivity is not as high as other metals), provides reproducible results and so is often used as a calibrating device. The main advantage of the thermistor over the platinum RTD is its high sensitivity and smaller size.

The following table lists the advantages, limitations and disadvantages of both types of temperature measuring instruments.

| Advantages of the Platinum RTD - The temperature v resistor relationship is linear within +0.4% over the temp. range -2000C to 400C. At 10000C the inaccuracy is only + 1.2% - The metal used for the coil (platinum) is very stable and provides reproducible results even though metals such as copper and nickel have a higher sensitivity. - The metal is chemically inert i.e., Unreactive and resists corrosion and oxidation. - Platinum has a high melting point of 17690C - The instrument can be used to measure low temperature differences and has most linear resistance-temperature characteristic. | Advantages of the Thermistor - The equipment is small in size and has a high sensitivity. - It is a relatively cheep method of detecting and measuring small changes in temperature. - It provides a fast output response to temperature changes because of its small size. - Its small size enables the temperature to be recorded at any point. - Small changes in temperature can cause large changes in resistances -gives accurate temperature measurement control. - Semi-conductor only have limited measurement range from -500C to 1500C inaccuracy is typically + 3% Limits their range of application |
| --- | --- |
| Limitations of the Platinum RTD - Nickel coil used mainly in the range -1000C to 3000C - The PRTD has greater accuracy when measuring temperature in the range -2000C to 8000C and best ratio price to performance for temperature range. - The upper range may be extended to 15000C if high melting materials used for the housing and coil construction - Measurement accuracy of + 0.5% or better. - Sensitivity approx.. 0.4 Ω per oC for 100 Ω | Limitations of Thermistor - The resistance of a typical Thermistor can vary from 400Ω at 00C to 100 Ω at 1400C - Temperature range is -1000C to 3000C - Gives large changes in resistor per degree changes in temperature giving accuracy + 0.10C but only possible over small range of temperature. |
| Disadvantages of the Platinum RTD - Platinum sensitivity is not as high as other metals such as copper and nickel - Expensive (cheaper resistors use platinum film deposited on ceramic substate - Slow response to temperature variation - This RTD tend to be fragile and easily damaged when subjected to excessive vibration or shock. | Disadvantages of Thermistor - Small sizes decrease its heat dissipation capability as self-heating effect becomes greater - Need to operate at low current levels when compared to the RTD - Have non linearity - Measurement range is limited - Measure and sensitivity are less than the RTD. |

Task 3 - Temperature Circuit Design

Having considered the data obtained during research (Task 1) and testing (Task 2) of the two sensors the platinum RTD sensor was selected. The selection was based on device range, accuracy, linearity reliability sensitivity and cost which are recorded in the table of advantages, limitations and disadvantages of each sensor in Task 2.

A measuring circuit to convert the resistance of the sensor to a comparative voltage was then designed for the given specification of the measuring device.

Specification:   Input range 0 0C to 100 0C and Output range 0 to 100 volts

Multi Sim Software was used to develop the circuit.

Figure 13. Screen shot of Multisim software package used to develop a signal conditioning circuit.

- The RTD sensor was connected to one arm of a Wheatstone bridge circuit, which converts its electrical resistance to a comparable output voltage.

- An amplifier (op amp) is included to amplify the bridge output voltage calibrated to give temperature.

- A spread sheet (Table 1) was created to obtain bridge output and amplifier output values for each of the platinum RTD (PT100) sensor resistance values listed.

- The resistance values were obtained from a standard reference chart for the temperature range 00C to 1000C in increments of 100C for a RTD (PT100) sensor.

| Temperature (0C ) | PT 100 ( Ω ) | Bridge Output ( mV) | Amplifier Output Voltage (V ) |
| --- | --- | --- | --- |
| 0 | 100 | 0.00 | 0.00 |
| 10 | 103.9 | 11.13 | 0.74 |
| 20 | 107.79 | 22.20 | 1.48 |
| 30 | 111.67 | 33.23 | 2.22 |
| 40 | 115.54 | 44.21 | 2.95 |
| 50 | 119.4 | 55.13 | 3.68 |
| 60 | 123.24 | 65.99 | 4.40 |
| 70 | 127.07 | 76.79 | 5.12 |
| 80 | 130.89 | 87.55 | 5.84 |
| 90 | 134.7 | 98.25 | 6.55 |
| 100 | 138.5 | 108.91 | 7.26 |

Table 1. Giving bridge output and Op Amp output voltages

The equations used to calculate the bridge and amplifier outputs are given below together with an example showing calculations for a platinum RTD (PT100) sensor at a temperature of 500C and a resistance of 119.4Ω.

Example

- Calculation for Bridge output

R2                                   PT100

Vin

R1                                    R3

- RTD (PT100) Temperature 50 0C from Table 1

- Resistance 119.4 Ω from Table 1

- R1 	= 4000 Ω  =  4 KΩ

- R2 	= 4000 Ω  =  4 KΩ

- PT100 =  119.4 Ω

- Vin     = 12v supply

Bridge out V0 (mV)

V0 =      PT100        -          R3            X   Vin

PT100 +R2          R3 + R1

=           119.4         -     100              X 12

119.4 + 4000     100 + 4000

=   [ 0.029 – 0.0244]  X 12

= 12 X 4.6 x 10-3

= 55.2 mV

Bridge output is 55.2 mV

- Calculation for amplifier output

R1 = Rin           = 1.5 KΩ

R2 = R feedback    = 100 KΩ

V1 = Vin          = 55.2 mV from (i) calcs.

Vo = Vout

Vout = Gain (voltage) x Vin

= R2    x Vin

R1

= 100   x 55.2 x 10-3

1.5

=  66.67 x 55.2 x 10-3

= 3.68 V

Amplifier output = 3.68 V

The above calculated values compared favourably with those listed in Table 1.

The two meters shown on the Multi Sim screen shot (figure 13) measure the output voltage of the Bridge and Op amp respectively for a RTD (PT100) sensor at 500C. On comparing the Multisim value with the corresponding results in Table 1 a discrepancy was noted. There is a difference of 4.35mV on the bridge output and 0.294V on the Op Amp output.

On investigation it was noted that Resistors R1 and R2 used on the bridge circuit should have read 4 KΩ not 3.95 KΩ and the inclusion of resistor R7 connecting the bridge output and Op amp non - inverting (+) terminal is a possible error and should have been omitted. These could possibly have caused the discrepancy noted.

Having developed the circuit on the Multisim package the circuit was built and tested. Unfortunately, it could not be amended due to lack of time and this facility was not available out of college hours because of the current “lockdown arrangements” regarding the COVID 19 pandemic. However, the assignment proved to be a valuable experience for future studies.

Task 4 - Signal Transmission

There is a need in many measuring systems to transmit measurement signals over quite large distances from the point of measurement to the place where the signals are recorded (Morris, 2006). This can create many problems for which a solution must be found. Of the many difficulties associated with long distance signal transmission, contamination of the measurement signal by noise is the most serious. Many sources of noise exist in industrial environments, such as electromagnetic fields from electrical machinery and power cables to spikes on the a.c power supply. Signals can be transmitted electrically, pneumatically, optically or by telemetry, in either analogue or digital format. There are various options for signal transmission ranging from electrical transmission to digital protocols transmission (Morris, 2006).

Digital transmission has very significant advantages compared with analogue transmission because the possibility of signal corruption during transmission is greatly reduced. The aim over many years has been to use intelligent devices to their full potential by making all communications, including the measured signal, entirely digital. Many different protocols exist for digital signal transmission such as Profibus being in widespread use. Due to the need to maintain compatibility with vast current investment in analogue instrumentation systems, a number of part analogue/part digital communication protocols have been developed as an interim measure. Prominent amongst these is a protocol called HART (Highway Addressable Remote Transducer) (Morris, 2006).

HART is a manufacturer – independent protocol that provides for analogue measurement signal transmission as well as sending command/status information digitally. In a process plant there are likely to be a large number of smart sensors, each providing information which has to be fed back to a control panel. To avoid using separate cables for each sensor to transmit their data, a bus system can be used. A bus is a common highway for signals which is used to link components. Each sensor would put its information onto the common highway for transmission to the control panel. The HART communication protocol is widely used for such transmissions (Bolton, 2015). This involves the digital signal from a smart sensor being superimposed on an analogue 4-20mA current loop signal. With this protocol, a 0 is represented by a 2200Hz frequency and a 1 by a 1200Hz frequency and these are superimposed on the d.c. signal to give simultaneous digital and analogue transmission (Petruzella, 2017). The digital data transfer rate is 1200 bits/s. The arrangement is that a master, such as a display terminal, sends a message with a request for data to a device, the device interprets the request and replies with the data (Bolton, 2015).

Figure 14. The HART communication protocol (Petruzella, 2017).

The HART system supports two modes of use, a hybrid mode and fully digital mode.

In hybrid mode status/command signals are digital but data transmission takes place in analogue form (usually 4-20mA format).  However, in this mode is not possible to transmit multiple measurement signals on a single bus since the analogue signals would corrupt each other. To overcome this the network must be arranged in a star configuration using a separate line for each field device rather than a common bus.

In fully digital mode data transmission is digital as well as status/command signals. This enables one cable to carry signals for up to fifteen intelligent devices (.  Using this mode, the data transmission speed is very low. Thus, the main application of the HART protocol has been to provide a communication capability with intelligent devices when existing analogue measurement signal transmission has to be retained because conversion to fully digital operation would be too expensive.

References

Petruzella, F. D. (2017). Programmable Logic Controllers. 5th ed. New York: Mc Graw Hill.

Morris, A. S. (2006). Measurement & Instrumentation Principles. 1st ed. Oxford: Elsevier Butterworth-Heinemann.

Bolton, W. (2015). Instrumentation & Control Systems. 2nd ed. Oxford: Newnes

Harrison, R. G. (2015). Meteorolgical Measurements and Instrumentation. 1st ed. London: John Wiley and Sons.

Bird, J. O. May, A. J. C. (1983). Engineering Science 3 Checkbook. 1st ed. London: Butterworth.

Appendix

| Time | RTD | IC | RTD Temp | IC Temp | % Error |
| --- | --- | --- | --- | --- | --- |
| Seconds | mV | V |  |  |  |
| 0 | 107.96 | 2.9839 | 20.67532 | 25.39 | 18.56902 |
| 1 | 108.11 | 2.985 | 21.06494 | 25.5 | 17.39241 |
| 2 | 108.02 | 2.9851 | 20.83117 | 25.51 | 18.34116 |
| 3 | 107.95 | 2.9855 | 20.64935 | 25.55 | 19.18062 |
| 4 | 107.98 | 2.9855 | 20.72727 | 25.55 | 18.87564 |
| 5 | 108.1 | 2.9863 | 21.03896 | 25.63 | 17.91275 |
| 6 | 107.98 | 2.9861 | 20.72727 | 25.61 | 19.06571 |
| 7 | 108.03 | 2.9861 | 20.85714 | 25.61 | 18.5586 |
| 8 | 108.12 | 2.9869 | 21.09091 | 25.69 | 17.90226 |
| 9 | 108.02 | 2.9846 | 20.83117 | 25.46 | 18.1808 |
| 10 | 108.08 | 2.9847 | 20.98701 | 25.47 | 17.60105 |
| 11 | 108.16 | 2.9851 | 21.19481 | 25.51 | 16.9157 |
| 12 | 108.13 | 2.9853 | 21.11688 | 25.53 | 17.286 |
| 13 | 108.31 | 2.9869 | 21.58442 | 25.69 | 15.98126 |
| 14 | 108.25 | 2.9864 | 21.42857 | 25.64 | 16.42523 |
| 15 | 108.14 | 2.9855 | 21.14286 | 25.55 | 17.24909 |
| 16 | 108.17 | 2.9865 | 21.22078 | 25.65 | 17.26792 |
| 17 | 108.38 | 2.9865 | 21.76623 | 25.65 | 15.14139 |
| 18 | 108.29 | 2.9867 | 21.53247 | 25.67 | 16.11816 |
| 19 | 108.32 | 2.9879 | 21.61039 | 25.79 | 16.20632 |
| 20 | 108.41 | 2.9868 | 21.84416 | 25.68 | 14.93709 |
| 21 | 108.41 | 2.9884 | 21.84416 | 25.84 | 15.46379 |
| 22 | 108.53 | 2.9892 | 22.15584 | 25.92 | 14.52221 |
| 23 | 108.54 | 2.9876 | 22.18182 | 25.76 | 13.89046 |
| 24 | 108.47 | 2.9885 | 22 | 25.85 | 14.89362 |
| 25 | 108.61 | 2.989 | 22.36364 | 25.9 | 13.65391 |
| 26 | 108.7 | 2.989 | 22.5974 | 25.9 | 12.75134 |
| 27 | 108.71 | 2.9891 | 22.62338 | 25.91 | 12.68477 |
| 28 | 108.64 | 2.9891 | 22.44156 | 25.91 | 13.3865 |
| 29 | 108.81 | 2.9887 | 22.88312 | 25.87 | 11.54574 |
| 30 | 108.98 | 2.9916 | 23.32468 | 26.16 | 10.8384 |
| 31 | 108.92 | 2.992 | 23.16883 | 26.2 | 11.56935 |
| 32 | 108.99 | 2.9925 | 23.35065 | 26.25 | 11.04515 |
| 33 | 108.98 | 2.9934 | 23.32468 | 26.34 | 11.4477 |
| 34 | 109.15 | 2.9933 | 23.76623 | 26.33 | 9.737054 |
| 35 | 109.27 | 2.9943 | 24.07792 | 26.43 | 8.899273 |
| 36 | 109.33 | 2.9947 | 24.23377 | 26.47 | 8.448182 |
| 37 | 109.36 | 2.9954 | 24.31169 | 26.54 | 8.39605 |
| 38 | 109.37 | 2.9953 | 24.33766 | 26.53 | 8.263617 |
| 39 | 109.5 | 2.9962 | 24.67532 | 26.62 | 7.305317 |
| 40 | 109.65 | 2.9966 | 25.06494 | 26.66 | 5.982989 |
| 41 | 109.51 | 2.9965 | 24.7013 | 26.65 | 7.3122 |
| 42 | 109.53 | 2.9965 | 24.75325 | 26.65 | 7.117273 |
| 43 | 109.72 | 2.9968 | 25.24675 | 26.68 | 5.371989 |
| 44 | 110.09 | 2.9971 | 26.20779 | 26.71 | 1.880224 |
| 45 | 110.02 | 2.9972 | 26.02597 | 26.72 | 2.597403 |
| 46 | 110.05 | 2.997 | 26.1039 | 26.7 | 2.232599 |
| 47 | 110.03 | 2.9972 | 26.05195 | 26.72 | 2.500194 |
| 48 | 110.15 | 2.9975 | 26.36364 | 26.75 | 1.44435 |
| 49 | 110.1 | 2.9975 | 26.23377 | 26.75 | 1.929846 |
| 50 | 110.31 | 2.9975 | 26.77922 | 26.75 | -0.10924 |
| 51 | 110.23 | 2.9975 | 26.57143 | 26.75 | 0.667557 |
| 52 | 110.23 | 2.9976 | 26.57143 | 26.76 | 0.704676 |
| 53 | 110.35 | 2.9976 | 26.88312 | 26.76 | -0.46008 |
| 54 | 110.51 | 2.9978 | 27.2987 | 26.78 | -1.9369 |
| 55 | 110.48 | 2.9978 | 27.22078 | 26.78 | -1.64593 |
| 56 | 110.49 | 2.9978 | 27.24675 | 26.78 | -1.74292 |
| 57 | 110.61 | 2.9978 | 27.55844 | 26.78 | -2.9068 |
| 58 | 110.74 | 2.9978 | 27.8961 | 26.78 | -4.16768 |
| 59 | 110.76 | 2.9978 | 27.94805 | 26.78 | -4.36166 |
| 60 | 110.87 | 2.9978 | 28.23377 | 26.78 | -5.42855 |
| 61 | 110.85 | 2.9978 | 28.18182 | 26.78 | -5.23457 |
| 62 | 110.9 | 2.9978 | 28.31169 | 26.78 | -5.71952 |
| 63 | 110.88 | 2.9978 | 28.25974 | 26.78 | -5.52554 |
| 64 | 110.89 | 2.9978 | 28.28571 | 26.78 | -5.62253 |
| 65 | 110.95 | 2.9978 | 28.44156 | 26.78 | -6.20448 |
| 66 | 110.99 | 2.9978 | 28.54545 | 26.78 | -6.59244 |
| 67 | 111.11 | 2.9978 | 28.85714 | 26.78 | -7.75632 |
| 68 | 111.12 | 2.9978 | 28.88312 | 26.78 | -7.85331 |
| 69 | 111.23 | 2.9978 | 29.16883 | 26.78 | -8.92021 |
| 70 | 111.24 | 2.9979 | 29.19481 | 26.79 | -8.9765 |
| 71 | 111.29 | 2.998 | 29.32468 | 26.8 | -9.42043 |
| 72 | 111.38 | 2.9979 | 29.55844 | 26.79 | -10.3339 |
| 73 | 111.37 | 2.998 | 29.53247 | 26.8 | -10.1958 |
| 74 | 111.4 | 2.9981 | 29.61039 | 26.81 | -10.4453 |
| 75 | 111.4 | 2.9982 | 29.61039 | 26.82 | -10.4041 |
| 76 | 111.39 | 2.9982 | 29.58442 | 26.82 | -10.3073 |
| 77 | 111.38 | 2.9983 | 29.55844 | 26.83 | -10.1694 |
| 78 | 111.5 | 2.9986 | 29.87013 | 26.86 | -11.2067 |
| 79 | 111.54 | 2.9987 | 29.97403 | 26.87 | -11.552 |
| 80 | 111.52 | 2.9989 | 29.92208 | 26.89 | -11.2759 |
| 81 | 111.58 | 2.9992 | 30.07792 | 26.92 | -11.7308 |
| 82 | 111.55 | 2.9993 | 30 | 26.93 | -11.3999 |
| 83 | 111.57 | 2.9997 | 30.05195 | 26.97 | -11.4273 |
| 84 | 111.6 | 3.0004 | 30.12987 | 27.04 | -11.427 |
| 85 | 111.61 | 3.0014 | 30.15584 | 27.14 | -11.1122 |
| 86 | 111.65 | 3.0013 | 30.25974 | 27.13 | -11.5361 |
| 87 | 111.66 | 3.0025 | 30.28571 | 27.25 | -11.1402 |
| 88 | 111.67 | 3.0041 | 30.31169 | 27.41 | -10.5862 |
| 89 | 111.64 | 3.0045 | 30.23377 | 27.45 | -10.1412 |
| 90 | 111.69 | 3.0071 | 30.36364 | 27.71 | -9.57646 |
| 91 | 111.71 | 3.0076 | 30.41558 | 27.76 | -9.56623 |
| 92 | 111.74 | 3.0087 | 30.49351 | 27.87 | -9.41337 |
| 93 | 111.71 | 3.0089 | 30.41558 | 27.89 | -9.05552 |
| 94 | 111.7 | 3.0113 | 30.38961 | 28.13 | -8.03274 |
| 95 | 111.73 | 3.0122 | 30.46753 | 28.22 | -7.96432 |
| 96 | 111.77 | 3.0133 | 30.57143 | 28.33 | -7.91186 |
| 97 | 111.78 | 3.0155 | 30.5974 | 28.55 | -7.17129 |
| 98 | 111.77 | 3.0175 | 30.57143 | 28.75 | -6.3354 |
| 99 | 111.81 | 3.0215 | 30.67532 | 29.15 | -5.23267 |
| 100 | 111.84 | 3.022 | 30.75325 | 29.2 | -5.31934 |
| 101 | 111.87 | 3.0225 | 30.83117 | 29.25 | -5.40571 |
| 102 | 111.88 | 3.0234 | 30.85714 | 29.34 | -5.1709 |
| 103 | 111.89 | 3.0254 | 30.88312 | 29.54 | -4.54677 |
| 104 | 111.91 | 3.0258 | 30.93506 | 29.58 | -4.58102 |
| 105 | 111.92 | 3.0282 | 30.96104 | 29.82 | -3.82642 |
| 106 | 111.95 | 3.0312 | 31.03896 | 30.12 | -3.051 |
| 107 | 112 | 3.0322 | 31.16883 | 30.22 | -3.13975 |
| 108 | 111.96 | 3.0332 | 31.06494 | 30.32 | -2.45691 |
| 109 | 112.08 | 3.0349 | 31.37662 | 30.49 | -2.90792 |
| 110 | 111.93 | 3.0335 | 30.98701 | 30.35 | -2.09889 |
| 111 | 112.08 | 3.0376 | 31.37662 | 30.76 | -2.00463 |
| 112 | 112.1 | 3.0391 | 31.42857 | 30.91 | -1.67768 |
| 113 | 112.24 | 3.0397 | 31.79221 | 30.97 | -2.65485 |
| 114 | 112.2 | 3.0404 | 31.68831 | 31.04 | -2.08863 |
| 115 | 112.26 | 3.0411 | 31.84416 | 31.11 | -2.35987 |
| 116 | 112.3 | 3.0419 | 31.94805 | 31.19 | -2.43043 |
| 117 | 112.3 | 3.0421 | 31.94805 | 31.21 | -2.36479 |
| 118 | 112.35 | 3.0423 | 32.07792 | 31.23 | -2.71509 |
| 119 | 112.33 | 3.0429 | 32.02597 | 31.29 | -2.35211 |
| 120 | 112.39 | 3.0435 | 32.18182 | 31.35 | -2.65333 |
| 121 | 112.5 | 3.0443 | 32.46753 | 31.43 | -3.30109 |
| 122 | 112.49 | 3.0456 | 32.44156 | 31.56 | -2.79328 |
| 123 | 112.54 | 3.0465 | 32.57143 | 31.65 | -2.91131 |
| 124 | 112.58 | 3.0472 | 32.67532 | 31.72 | -3.01174 |
| 125 | 112.61 | 3.0488 | 32.75325 | 31.88 | -2.73917 |
| 126 | 112.76 | 3.05 | 33.14286 | 32 | -3.57143 |
| 127 | 112.69 | 3.0519 | 32.96104 | 32.19 | -2.39527 |
| 128 | 112.74 | 3.0525 | 33.09091 | 32.25 | -2.60747 |
| 129 | 112.81 | 3.0542 | 33.27273 | 32.42 | -2.63025 |
| 130 | 112.86 | 3.0583 | 33.4026 | 32.83 | -1.74413 |
| 131 | 112.91 | 3.0589 | 33.53247 | 32.89 | -1.95338 |
| 132 | 113.02 | 3.0597 | 33.81818 | 32.97 | -2.57259 |
| 133 | 113.12 | 3.063 | 34.07792 | 33.3 | -2.3361 |
| 134 | 112.99 | 3.0656 | 33.74026 | 33.56 | -0.53713 |
| 135 | 112.97 | 3.0656 | 33.68831 | 33.56 | -0.38234 |
| 136 | 113.09 | 3.0693 | 34 | 33.93 | -0.20631 |
| 137 | 113.26 | 3.0727 | 34.44156 | 34.27 | -0.50061 |
| 138 | 113.3 | 3.0728 | 34.54545 | 34.28 | -0.77437 |
| 139 | 113.32 | 3.0734 | 34.5974 | 34.34 | -0.74957 |
| 140 | 113.46 | 3.0762 | 34.96104 | 34.62 | -0.98509 |
| 141 | 113.51 | 3.0784 | 35.09091 | 34.84 | -0.72018 |
| 142 | 113.54 | 3.0779 | 35.16883 | 34.79 | -1.08891 |
| 143 | 113.72 | 3.0799 | 35.63636 | 34.99 | -1.84728 |
| 144 | 113.67 | 3.0811 | 35.50649 | 35.11 | -1.12929 |
| 145 | 113.89 | 3.0825 | 36.07792 | 35.25 | -2.34872 |
| 146 | 114 | 3.0831 | 36.36364 | 35.31 | -2.98396 |
| 147 | 114.05 | 3.0833 | 36.49351 | 35.33 | -3.29325 |
| 148 | 114.07 | 3.0841 | 36.54545 | 35.41 | -3.20659 |
| 149 | 114.26 | 3.0844 | 37.03896 | 35.44 | -4.51174 |
| 150 | 114.35 | 3.0847 | 37.27273 | 35.47 | -5.0824 |
| 151 | 114.36 | 3.0856 | 37.2987 | 35.56 | -4.88949 |
| 152 | 114.39 | 3.0858 | 37.37662 | 35.58 | -5.04953 |
| 153 | 114.55 | 3.086 | 37.79221 | 35.6 | -6.15789 |
| 154 | 114.73 | 3.0863 | 38.25974 | 35.63 | -7.38069 |
| 155 | 114.75 | 3.0865 | 38.31169 | 35.65 | -7.46617 |
| 156 | 114.6 | 3.0866 | 37.92208 | 35.66 | -6.34346 |
| 157 | 114.72 | 3.0867 | 38.23377 | 35.67 | -7.18746 |
| 158 | 114.79 | 3.0868 | 38.41558 | 35.68 | -7.667 |
| 159 | 114.9 | 3.0868 | 38.7013 | 35.68 | -8.46777 |
| 160 | 114.96 | 3.0869 | 38.85714 | 35.69 | -8.87403 |
| 161 | 114.9 | 3.087 | 38.7013 | 35.7 | -8.407 |
| 162 | 114.9 | 3.0871 | 38.7013 | 35.71 | -8.37664 |
| 163 | 114.89 | 3.0871 | 38.67532 | 35.71 | -8.30391 |
| 164 | 115.02 | 3.0873 | 39.01299 | 35.73 | -9.18832 |
| 165 | 115.21 | 3.0874 | 39.50649 | 35.74 | -10.5386 |
| 166 | 115.21 | 3.0874 | 39.50649 | 35.74 | -10.5386 |
| 167 | 115.24 | 3.0876 | 39.58442 | 35.76 | -10.6947 |
| 168 | 115.22 | 3.0878 | 39.53247 | 35.78 | -10.4876 |
| 169 | 115.38 | 3.0881 | 39.94805 | 35.81 | -11.5556 |
| 170 | 115.46 | 3.0879 | 40.15584 | 35.79 | -12.1985 |
| 171 | 115.4 | 3.0881 | 40 | 35.81 | -11.7006 |
| 172 | 115.51 | 3.0884 | 40.28571 | 35.84 | -12.4043 |
| 173 | 115.63 | 3.0884 | 40.5974 | 35.84 | -13.274 |
| 174 | 115.65 | 3.0893 | 40.64935 | 35.93 | -13.1348 |
| 175 | 115.6 | 3.0896 | 40.51948 | 35.96 | -12.6793 |
| 176 | 115.73 | 3.0899 | 40.85714 | 35.99 | -13.5236 |
| 177 | 115.69 | 3.0911 | 40.75325 | 36.11 | -12.8586 |
| 178 | 115.74 | 3.092 | 40.88312 | 36.2 | -12.9368 |
| 179 | 115.75 | 3.0924 | 40.90909 | 36.24 | -12.8838 |
| 180 | 115.76 | 3.0924 | 40.93506 | 36.24 | -12.9555 |
| 181 | 115.81 | 3.0941 | 41.06494 | 36.41 | -12.7848 |
| 182 | 115.77 | 3.0954 | 40.96104 | 36.54 | -12.0992 |
| 183 | 115.88 | 3.0965 | 41.24675 | 36.65 | -12.5423 |
| 184 | 115.87 | 3.0979 | 41.22078 | 36.79 | -12.0434 |
| 185 | 115.97 | 3.0977 | 41.48052 | 36.77 | -12.8108 |
| 186 | 115.86 | 3.1005 | 41.19481 | 37.05 | -11.1871 |
| 187 | 115.98 | 3.1009 | 41.50649 | 37.09 | -11.9075 |
| 188 | 115.97 | 3.103 | 41.48052 | 37.3 | -11.2078 |
| 189 | 115.96 | 3.1035 | 41.45455 | 37.35 | -10.9894 |
| 190 | 115.99 | 3.1061 | 41.53247 | 37.61 | -10.4293 |
| 191 | 116.02 | 3.1059 | 41.61039 | 37.59 | -10.6954 |
| 192 | 115.97 | 3.109 | 41.48052 | 37.9 | -9.44728 |
| 193 | 115.99 | 3.1108 | 41.53247 | 38.08 | -9.06635 |
| 194 | 116.01 | 3.1134 | 41.58442 | 38.34 | -8.46222 |
| 195 | 116.04 | 3.113 | 41.66234 | 38.3 | -8.77895 |
| 196 | 116.02 | 3.1154 | 41.61039 | 38.54 | -7.96676 |
| 197 | 116.08 | 3.1181 | 41.76623 | 38.81 | -7.6172 |
| 198 | 116.09 | 3.119 | 41.79221 | 38.9 | -7.43498 |
| 199 | 116.1 | 3.1208 | 41.81818 | 39.08 | -7.00661 |
| 200 | 116.15 | 3.1215 | 41.94805 | 39.15 | -7.147 |
| 201 | 116.12 | 3.1234 | 41.87013 | 39.34 | -6.43144 |
| 202 | 116.13 | 3.1244 | 41.8961 | 39.44 | -6.22744 |
| 203 | 116.13 | 3.1242 | 41.8961 | 39.42 | -6.28134 |
| 204 | 116.12 | 3.125 | 41.87013 | 39.5 | -6.00033 |
| 205 | 116.14 | 3.1253 | 41.92208 | 39.53 | -6.0513 |
| 206 | 116.17 | 3.1257 | 42 | 39.57 | -6.14102 |
| 207 | 116.18 | 3.1266 | 42.02597 | 39.66 | -5.96564 |
| 208 | 116.18 | 3.127 | 42.02597 | 39.7 | -5.85888 |
| 209 | 116.2 | 3.1277 | 42.07792 | 39.77 | -5.80317 |
| 210 | 116.2 | 3.1285 | 42.07792 | 39.85 | -5.59077 |
| 211 | 116.2 | 3.1296 | 42.07792 | 39.96 | -5.30011 |
| 212 | 116.24 | 3.1299 | 42.18182 | 39.99 | -5.48092 |
| 213 | 116.22 | 3.13 | 42.12987 | 40 | -5.32468 |
| 214 | 116.22 | 3.1303 | 42.12987 | 40.03 | -5.24574 |
| 215 | 116.25 | 3.1309 | 42.20779 | 40.09 | -5.28259 |
| 216 | 116.26 | 3.1312 | 42.23377 | 40.12 | -5.26861 |
| 217 | 116.26 | 3.1312 | 42.23377 | 40.12 | -5.26861 |
| 218 | 116.29 | 3.1314 | 42.31169 | 40.14 | -5.41028 |
| 219 | 116.29 | 3.132 | 42.31169 | 40.2 | -5.25296 |
| 220 | 116.32 | 3.1319 | 42.38961 | 40.19 | -5.47303 |
| 221 | 116.31 | 3.1322 | 42.36364 | 40.22 | -5.32978 |
| 222 | 116.35 | 3.1323 | 42.46753 | 40.23 | -5.56185 |
| 223 | 116.33 | 3.1324 | 42.41558 | 40.24 | -5.40652 |
| 224 | 116.32 | 3.1328 | 42.38961 | 40.28 | -5.23736 |
| 225 | 116.36 | 3.1329 | 42.49351 | 40.29 | -5.46912 |
| 226 | 116.33 | 3.1335 | 42.41558 | 40.35 | -5.11917 |
| 227 | 116.37 | 3.1342 | 42.51948 | 40.42 | -5.19416 |
| 228 | 116.41 | 3.1349 | 42.62338 | 40.49 | -5.2689 |
| 229 | 116.39 | 3.1348 | 42.57143 | 40.48 | -5.16657 |
| 230 | 116.42 | 3.1362 | 42.64935 | 40.62 | -4.99594 |
| 231 | 116.41 | 3.1367 | 42.62338 | 40.67 | -4.80299 |
| 232 | 116.39 | 3.1386 | 42.57143 | 40.86 | -4.18852 |
| 233 | 116.43 | 3.1388 | 42.67532 | 40.88 | -4.39169 |
| 234 | 116.44 | 3.14 | 42.7013 | 41 | -4.14951 |
| 235 | 116.47 | 3.141 | 42.77922 | 41.1 | -4.0857 |
| 236 | 116.51 | 3.1411 | 42.88312 | 41.11 | -4.3131 |
| 237 | 116.55 | 3.1416 | 42.98701 | 41.16 | -4.43881 |
| 238 | 116.48 | 3.1448 | 42.80519 | 41.48 | -3.19478 |
| 239 | 116.46 | 3.147 | 42.75325 | 41.7 | -2.52577 |
| 240 | 116.53 | 3.15 | 42.93506 | 42 | -2.22635 |
| 241 | 116.53 | 3.1482 | 42.93506 | 41.82 | -2.66634 |
| 242 | 116.58 | 3.15 | 43.06494 | 42 | -2.53556 |
| 243 | 116.62 | 3.1541 | 43.16883 | 42.41 | -1.78927 |
| 244 | 116.6 | 3.1551 | 43.11688 | 42.51 | -1.42762 |
| 245 | 116.61 | 3.1569 | 43.14286 | 42.69 | -1.0608 |
| 246 | 116.62 | 3.1554 | 43.16883 | 42.54 | -1.47821 |
| 247 | 116.68 | 3.1568 | 43.32468 | 42.68 | -1.51049 |
| 248 | 116.69 | 3.1618 | 43.35065 | 43.18 | -0.3952 |
| 249 | 116.71 | 3.1612 | 43.4026 | 43.12 | -0.65537 |
| 250 | 116.81 | 3.1643 | 43.66234 | 43.43 | -0.53497 |
| 251 | 116.74 | 3.1649 | 43.48052 | 43.49 | 0.021799 |
| 252 | 116.78 | 3.1655 | 43.58442 | 43.55 | -0.07903 |
| 253 | 116.82 | 3.1682 | 43.68831 | 43.82 | 0.300521 |
| 254 | 116.93 | 3.1691 | 43.97403 | 43.91 | -0.14581 |
| 255 | 116.95 | 3.1694 | 44.02597 | 43.94 | -0.19566 |
| 256 | 116.89 | 3.1696 | 43.87013 | 43.96 | 0.204436 |
| 257 | 116.88 | 3.1719 | 43.84416 | 44.19 | 0.78263 |
| 258 | 117 | 3.1723 | 44.15584 | 44.23 | 0.16766 |
| 259 | 116.98 | 3.1734 | 44.1039 | 44.34 | 0.532485 |
| 260 | 117.09 | 3.1731 | 44.38961 | 44.31 | -0.17967 |
| 261 | 117.14 | 3.1736 | 44.51948 | 44.36 | -0.35951 |
| 262 | 117.15 | 3.1738 | 44.54545 | 44.38 | -0.37281 |
| 263 | 117.21 | 3.1749 | 44.7013 | 44.49 | -0.47494 |
| 264 | 117.18 | 3.1749 | 44.62338 | 44.49 | -0.29979 |
| 265 | 117.33 | 3.175 | 45.01299 | 44.5 | -1.15278 |
| 266 | 117.49 | 3.1751 | 45.42857 | 44.51 | -2.06374 |
| 267 | 117.54 | 3.1753 | 45.55844 | 44.53 | -2.30955 |
| 268 | 117.49 | 3.1755 | 45.42857 | 44.55 | -1.9721 |
| 269 | 117.58 | 3.1758 | 45.66234 | 44.58 | -2.42785 |
| 270 | 117.53 | 3.1758 | 45.53247 | 44.58 | -2.13654 |
| 271 | 117.54 | 3.1758 | 45.55844 | 44.58 | -2.1948 |
| 272 | 117.64 | 3.1759 | 45.81818 | 44.59 | -2.75439 |
| 273 | 117.82 | 3.176 | 46.28571 | 44.6 | -3.77963 |
| 274 | 117.82 | 3.1762 | 46.28571 | 44.62 | -3.73311 |
| 275 | 117.75 | 3.1761 | 46.1039 | 44.61 | -3.34879 |
| 276 | 117.9 | 3.1761 | 46.49351 | 44.61 | -4.22216 |
| 277 | 117.87 | 3.1762 | 46.41558 | 44.62 | -4.02417 |
| 278 | 117.95 | 3.1762 | 46.62338 | 44.62 | -4.48986 |
| 279 | 117.94 | 3.1762 | 46.5974 | 44.62 | -4.43165 |
| 280 | 118.07 | 3.1762 | 46.93506 | 44.62 | -5.1884 |
| 281 | 118.06 | 3.1762 | 46.90909 | 44.62 | -5.13019 |
| 282 | 118.22 | 3.1762 | 47.32468 | 44.62 | -6.06158 |
| 283 | 118.16 | 3.1762 | 47.16883 | 44.62 | -5.71231 |
| 284 | 118.24 | 3.1762 | 47.37662 | 44.62 | -6.178 |
| 285 | 118.48 | 3.1762 | 48 | 44.62 | -7.57508 |
| 286 | 118.53 | 3.1762 | 48.12987 | 44.62 | -7.86614 |
| 287 | 118.61 | 3.1762 | 48.33766 | 44.62 | -8.33183 |
| 288 | 118.49 | 3.1762 | 48.02597 | 44.62 | -7.63329 |
| 289 | 118.51 | 3.1762 | 48.07792 | 44.62 | -7.74971 |
| 290 | 118.61 | 3.1762 | 48.33766 | 44.62 | -8.33183 |
| 291 | 118.75 | 3.1763 | 48.7013 | 44.63 | -9.12234 |
| 292 | 118.86 | 3.1762 | 48.98701 | 44.62 | -9.78712 |
| 293 | 118.81 | 3.1763 | 48.85714 | 44.63 | -9.47153 |
| 294 | 118.86 | 3.1763 | 48.98701 | 44.63 | -9.76252 |
| 295 | 118.92 | 3.1764 | 49.14286 | 44.64 | -10.087 |
| 296 | 118.98 | 3.1764 | 49.2987 | 44.64 | -10.4362 |
| 297 | 119.03 | 3.1764 | 49.42857 | 44.64 | -10.7271 |
| 298 | 118.9 | 3.1765 | 49.09091 | 44.65 | -9.94604 |
| 299 | 119.03 | 3.1765 | 49.42857 | 44.65 | -10.7023 |
| 300 | 119.17 | 3.1766 | 49.79221 | 44.66 | -11.4917 |
| 301 | 119.01 | 3.1766 | 49.37662 | 44.66 | -10.5612 |
| 302 | 119 | 3.1765 | 49.35065 | 44.65 | -10.5278 |
| 303 | 119.08 | 3.1768 | 49.55844 | 44.68 | -10.9186 |
| 304 | 119.01 | 3.1768 | 49.37662 | 44.68 | -10.5117 |
| 305 | 119.06 | 3.1768 | 49.50649 | 44.68 | -10.8024 |
| 306 | 119.19 | 3.1772 | 49.84416 | 44.72 | -11.4583 |
| 307 | 119.28 | 3.1779 | 50.07792 | 44.79 | -11.806 |
| 308 | 119.15 | 3.1776 | 49.74026 | 44.76 | -11.1266 |
| 309 | 119.41 | 3.1776 | 50.41558 | 44.76 | -12.6354 |
| 310 | 119.45 | 3.1776 | 50.51948 | 44.76 | -12.8675 |
| 311 | 119.42 | 3.1778 | 50.44156 | 44.78 | -12.6431 |
| 312 | 119.37 | 3.1782 | 50.31169 | 44.82 | -12.2528 |
| 313 | 119.42 | 3.1784 | 50.44156 | 44.84 | -12.4923 |
| 314 | 119.62 | 3.1784 | 50.96104 | 44.84 | -13.6508 |
| 315 | 119.67 | 3.1788 | 51.09091 | 44.88 | -13.8389 |
| 316 | 119.6 | 3.1793 | 50.90909 | 44.93 | -13.3076 |
| 317 | 119.72 | 3.1797 | 51.22078 | 44.97 | -13.8999 |
| 318 | 119.67 | 3.1797 | 51.09091 | 44.97 | -13.6111 |
| 319 | 119.69 | 3.1804 | 51.14286 | 45.04 | -13.5499 |
| 320 | 119.82 | 3.181 | 51.48052 | 45.1 | -14.1475 |
| 321 | 119.83 | 3.1817 | 51.50649 | 45.17 | -14.0281 |
| 322 | 119.82 | 3.1823 | 51.48052 | 45.23 | -13.8194 |
| 323 | 119.85 | 3.1833 | 51.55844 | 45.33 | -13.7402 |
| 324 | 119.97 | 3.1835 | 51.87013 | 45.35 | -14.3774 |
| 325 | 119.88 | 3.1846 | 51.63636 | 45.46 | -13.5864 |
| 326 | 119.91 | 3.1861 | 51.71429 | 45.61 | -13.3837 |
| 327 | 119.95 | 3.1852 | 51.81818 | 45.52 | -13.8361 |
| 328 | 119.99 | 3.1874 | 51.92208 | 45.74 | -13.5157 |
| 329 | 119.91 | 3.1889 | 51.71429 | 45.89 | -12.6918 |
| 330 | 119.85 | 3.1896 | 51.55844 | 45.96 | -12.1811 |
| 331 | 119.89 | 3.1922 | 51.66234 | 46.22 | -11.7749 |
| 332 | 119.9 | 3.1933 | 51.68831 | 46.33 | -11.5655 |
| 333 | 119.94 | 3.1947 | 51.79221 | 46.47 | -11.453 |
| 334 | 119.98 | 3.1955 | 51.8961 | 46.55 | -11.4846 |
| 335 | 119.99 | 3.1945 | 51.92208 | 46.45 | -11.7806 |
| 336 | 120.03 | 3.1959 | 52.02597 | 46.59 | -11.6677 |
| 337 | 120.1 | 3.1979 | 52.20779 | 46.79 | -11.579 |
| 338 | 120.15 | 3.1995 | 52.33766 | 46.95 | -11.4753 |
| 339 | 120.17 | 3.1998 | 52.38961 | 46.98 | -11.5147 |
| 340 | 120.15 | 3.1998 | 52.33766 | 46.98 | -11.4041 |
| 341 | 120.2 | 3.2013 | 52.46753 | 47.13 | -11.3251 |
| 342 | 120.15 | 3.2014 | 52.33766 | 47.14 | -11.026 |
| 343 | 120.22 | 3.2028 | 52.51948 | 47.28 | -11.0818 |
| 344 | 120.17 | 3.2035 | 52.38961 | 47.35 | -10.6433 |
| 345 | 120.25 | 3.2042 | 52.5974 | 47.42 | -10.9182 |
| 346 | 120.26 | 3.2059 | 52.62338 | 47.59 | -10.5765 |
| 347 | 120.32 | 3.2073 | 52.77922 | 47.73 | -10.5787 |
| 348 | 120.26 | 3.2086 | 52.62338 | 47.86 | -9.95273 |
| 349 | 120.4 | 3.2096 | 52.98701 | 47.96 | -10.4817 |
| 350 | 120.4 | 3.2099 | 52.98701 | 47.99 | -10.4126 |
| 351 | 120.41 | 3.2112 | 53.01299 | 48.12 | -10.1683 |
| 352 | 120.45 | 3.2115 | 53.11688 | 48.15 | -10.3154 |
| 353 | 120.42 | 3.2124 | 53.03896 | 48.24 | -9.9481 |
| 354 | 120.49 | 3.2123 | 53.22078 | 48.23 | -10.3479 |
| 355 | 120.4 | 3.213 | 52.98701 | 48.3 | -9.70396 |
| 356 | 120.46 | 3.2142 | 53.14286 | 48.42 | -9.75394 |
| 357 | 120.44 | 3.2153 | 53.09091 | 48.53 | -9.39812 |
| 358 | 120.43 | 3.2154 | 53.06494 | 48.54 | -9.32207 |
| 359 | 120.47 | 3.2152 | 53.16883 | 48.52 | -9.58127 |
| 360 | 120.46 | 3.2167 | 53.14286 | 48.67 | -9.19017 |
| 361 | 120.47 | 3.2166 | 53.16883 | 48.66 | -9.26599 |
| 362 | 120.46 | 3.2173 | 53.14286 | 48.73 | -9.05573 |
| 363 | 120.48 | 3.218 | 53.19481 | 48.8 | -9.00575 |
| 364 | 120.49 | 3.2183 | 53.22078 | 48.83 | -8.99197 |
| 365 | 120.51 | 3.218 | 53.27273 | 48.8 | -9.16542 |
| 366 | 120.52 | 3.2184 | 53.2987 | 48.84 | -9.1292 |
| 367 | 120.5 | 3.2186 | 53.24675 | 48.86 | -8.97821 |
| 368 | 120.52 | 3.219 | 53.2987 | 48.9 | -8.9953 |
| 369 | 120.51 | 3.2195 | 53.27273 | 48.95 | -8.8309 |
| 370 | 120.53 | 3.2196 | 53.32468 | 48.96 | -8.91478 |
| 371 | 120.55 | 3.2203 | 53.37662 | 49.03 | -8.86523 |
| 372 | 120.55 | 3.2204 | 53.37662 | 49.04 | -8.84303 |
| 373 | 120.59 | 3.2205 | 53.48052 | 49.05 | -9.03266 |
| 374 | 120.6 | 3.2204 | 53.50649 | 49.04 | -9.10786 |
| 375 | 120.6 | 3.2208 | 53.50649 | 49.08 | -9.01894 |
| 376 | 120.59 | 3.2208 | 53.48052 | 49.08 | -8.96601 |
| 377 | 120.64 | 3.221 | 53.61039 | 49.1 | -9.18613 |
| 378 | 120.61 | 3.2211 | 53.53247 | 49.11 | -9.00523 |
| 379 | 120.62 | 3.2216 | 53.55844 | 49.16 | -8.9472 |
| 380 | 120.64 | 3.2218 | 53.61039 | 49.18 | -9.00852 |
| 381 | 120.64 | 3.2219 | 53.61039 | 49.19 | -8.98636 |
| 382 | 120.64 | 3.2221 | 53.61039 | 49.21 | -8.94206 |
| 383 | 120.64 | 3.2222 | 53.61039 | 49.22 | -8.91993 |
| 384 | 120.64 | 3.2222 | 53.61039 | 49.22 | -8.91993 |
| 385 | 120.66 | 3.223 | 53.66234 | 49.3 | -8.84856 |
| 386 | 120.63 | 3.223 | 53.58442 | 49.3 | -8.6905 |
| 387 | 120.63 | 3.2236 | 53.58442 | 49.36 | -8.55838 |
| 388 | 120.63 | 3.2242 | 53.58442 | 49.42 | -8.42658 |
| 389 | 120.63 | 3.2241 | 53.58442 | 49.41 | -8.44852 |
| 390 | 120.63 | 3.2238 | 53.58442 | 49.38 | -8.51441 |
| 391 | 120.64 | 3.2248 | 53.61039 | 49.48 | -8.34759 |
| 392 | 120.66 | 3.2254 | 53.66234 | 49.54 | -8.32123 |
| 393 | 120.65 | 3.2267 | 53.63636 | 49.67 | -7.98543 |
| 394 | 120.68 | 3.2278 | 53.71429 | 49.78 | -7.90335 |
| 395 | 120.69 | 3.2276 | 53.74026 | 49.76 | -7.99891 |
| 396 | 120.69 | 3.2271 | 53.74026 | 49.71 | -8.10754 |
| 397 | 120.68 | 3.2289 | 53.71429 | 49.89 | -7.66544 |
| 398 | 120.71 | 3.2302 | 53.79221 | 50.02 | -7.5414 |
| 399 | 120.72 | 3.2314 | 53.81818 | 50.14 | -7.33582 |
| 400 | 120.71 | 3.231 | 53.79221 | 50.1 | -7.36968 |
| 401 | 120.72 | 3.2319 | 53.81818 | 50.19 | -7.22889 |
| 402 | 120.71 | 3.2322 | 53.79221 | 50.22 | -7.11312 |
| 403 | 120.72 | 3.2329 | 53.81818 | 50.29 | -7.01567 |
| 404 | 120.7 | 3.2349 | 53.76623 | 50.49 | -6.48888 |
| 405 | 120.72 | 3.2379 | 53.81818 | 50.79 | -5.96216 |
| 406 | 120.7 | 3.2371 | 53.76623 | 50.71 | -6.02689 |
| 407 | 120.69 | 3.2385 | 53.74026 | 50.85 | -5.68389 |
| 408 | 120.71 | 3.2413 | 53.79221 | 51.13 | -5.20674 |
| 409 | 120.71 | 3.2405 | 53.79221 | 51.05 | -5.37161 |
| 410 | 120.69 | 3.2415 | 53.74026 | 51.15 | -5.06405 |
| 411 | 120.73 | 3.2412 | 53.84416 | 51.12 | -5.32894 |
| 412 | 120.7 | 3.2435 | 53.76623 | 51.35 | -4.70542 |
| 413 | 120.71 | 3.244 | 53.79221 | 51.4 | -4.6541 |
| 414 | 120.76 | 3.2457 | 53.92208 | 51.57 | -4.56094 |
| 415 | 120.73 | 3.2439 | 53.84416 | 51.39 | -4.77555 |
| 416 | 120.73 | 3.2437 | 53.84416 | 51.37 | -4.81634 |
| 417 | 120.72 | 3.2453 | 53.81818 | 51.53 | -4.44048 |
| 418 | 120.73 | 3.2464 | 53.84416 | 51.64 | -4.26831 |
| 419 | 120.75 | 3.2469 | 53.8961 | 51.69 | -4.26795 |
| 420 | 120.74 | 3.2505 | 53.87013 | 52.05 | -3.49689 |
| 421 | 120.73 | 3.2501 | 53.84416 | 52.01 | -3.52654 |
| 422 | 120.75 | 3.25 | 53.8961 | 52 | -3.64635 |
| 423 | 120.76 | 3.2506 | 53.92208 | 52.06 | -3.57679 |
| 424 | 120.73 | 3.2516 | 53.84416 | 52.16 | -3.22883 |
| 425 | 120.73 | 3.2514 | 53.84416 | 52.14 | -3.26842 |
| 426 | 120.75 | 3.2538 | 53.8961 | 52.38 | -2.89443 |
| 427 | 120.74 | 3.254 | 53.87013 | 52.4 | -2.80559 |
| 428 | 120.75 | 3.255 | 53.8961 | 52.5 | -2.65925 |
| 429 | 120.76 | 3.2561 | 53.92208 | 52.61 | -2.49397 |
| 430 | 120.75 | 3.2573 | 53.8961 | 52.73 | -2.21146 |
| 431 | 120.75 | 3.2577 | 53.8961 | 52.77 | -2.13399 |
| 432 | 120.77 | 3.2593 | 53.94805 | 52.93 | -1.92339 |
| 433 | 120.75 | 3.2579 | 53.8961 | 52.79 | -2.09529 |
| 434 | 120.76 | 3.2587 | 53.92208 | 52.87 | -1.98993 |
| 435 | 120.77 | 3.259 | 53.94805 | 52.9 | -1.98119 |
| 436 | 120.76 | 3.2585 | 53.92208 | 52.85 | -2.02853 |
| 437 | 120.77 | 3.2597 | 53.94805 | 52.97 | -1.84643 |
| 438 | 120.75 | 3.2605 | 53.8961 | 53.05 | -1.59492 |
| 439 | 120.77 | 3.261 | 53.94805 | 53.1 | -1.59708 |
| 440 | 120.76 | 3.2617 | 53.92208 | 53.17 | -1.41448 |
| 441 | 120.78 | 3.2619 | 53.97403 | 53.19 | -1.47401 |
| 442 | 120.76 | 3.2621 | 53.92208 | 53.21 | -1.33824 |
| 443 | 120.77 | 3.263 | 53.94805 | 53.3 | -1.21586 |
| 444 | 120.79 | 3.2632 | 54 | 53.32 | -1.27532 |
| 445 | 120.76 | 3.2631 | 53.92208 | 53.31 | -1.14815 |
| 446 | 120.77 | 3.2626 | 53.94805 | 53.26 | -1.29187 |
| 447 | 120.79 | 3.2633 | 54 | 53.33 | -1.25633 |
| 448 | 120.78 | 3.2634 | 53.97403 | 53.34 | -1.18865 |
| 449 | 120.79 | 3.2636 | 54 | 53.36 | -1.1994 |
| 450 | 120.79 | 3.2639 | 54 | 53.39 | -1.14254 |
| 451 | 120.78 | 3.2642 | 53.97403 | 53.42 | -1.03711 |
| 452 | 120.78 | 3.2642 | 53.97403 | 53.42 | -1.03711 |
| 453 | 120.79 | 3.2644 | 54 | 53.44 | -1.0479 |
| 454 | 120.79 | 3.2646 | 54 | 53.46 | -1.0101 |
| 455 | 120.77 | 3.2647 | 53.94805 | 53.47 | -0.89406 |
| 456 | 120.79 | 3.2647 | 54 | 53.47 | -0.99121 |
| 457 | 120.79 | 3.2648 | 54 | 53.48 | -0.97233 |
| 458 | 120.79 | 3.2648 | 54 | 53.48 | -0.97233 |
| 459 | 120.79 | 3.2649 | 54 | 53.49 | -0.95345 |
| 460 | 120.76 | 3.2651 | 53.92208 | 53.51 | -0.7701 |
| 461 | 120.78 | 3.2651 | 53.97403 | 53.51 | -0.86718 |
| 462 | 120.8 | 3.2651 | 54.02597 | 53.51 | -0.96426 |
| 463 | 120.8 | 3.2652 | 54.02597 | 53.52 | -0.94539 |
| 464 | 120.79 | 3.2651 | 54 | 53.51 | -0.91572 |
| 465 | 120.78 | 3.2651 | 53.97403 | 53.51 | -0.86718 |
| 466 | 120.79 | 3.2653 | 54 | 53.53 | -0.87801 |
| 467 | 120.79 | 3.2653 | 54 | 53.53 | -0.87801 |
| 468 | 120.79 | 3.2653 | 54 | 53.53 | -0.87801 |
| 469 | 120.8 | 3.2652 | 54.02597 | 53.52 | -0.94539 |
| 470 | 120.8 | 3.2653 | 54.02597 | 53.53 | -0.92653 |
| 471 | 120.8 | 3.2653 | 54.02597 | 53.53 | -0.92653 |
| 472 | 120.79 | 3.2653 | 54 | 53.53 | -0.87801 |
| 473 | 120.8 | 3.2654 | 54.02597 | 53.54 | -0.90768 |
| 474 | 120.81 | 3.2654 | 54.05195 | 53.54 | -0.9562 |
| 475 | 120.78 | 3.2654 | 53.97403 | 53.54 | -0.81066 |
| 476 | 120.8 | 3.2654 | 54.02597 | 53.54 | -0.90768 |
| 477 | 120.81 | 3.2654 | 54.05195 | 53.54 | -0.9562 |
| 478 | 120.8 | 3.2654 | 54.02597 | 53.54 | -0.90768 |
| 479 | 120.81 | 3.2654 | 54.05195 | 53.54 | -0.9562 |
| 480 | 120.8 | 3.2654 | 54.02597 | 53.54 | -0.90768 |
| 481 | 120.8 | 3.2654 | 54.02597 | 53.54 | -0.90768 |
| 482 | 120.82 | 3.2655 | 54.07792 | 53.55 | -0.98585 |
| 483 | 120.8 | 3.2655 | 54.02597 | 53.55 | -0.88884 |
| 484 | 120.79 | 3.2654 | 54 | 53.54 | -0.85917 |
| 485 | 120.81 | 3.2656 | 54.05195 | 53.56 | -0.9185 |
| 486 | 120.82 | 3.2655 | 54.07792 | 53.55 | -0.98585 |
| 487 | 120.8 | 3.2655 | 54.02597 | 53.55 | -0.88884 |
| 488 | 120.82 | 3.2656 | 54.07792 | 53.56 | -0.96699 |
| 489 | 120.81 | 3.2657 | 54.05195 | 53.57 | -0.89966 |
| 490 | 120.82 | 3.2657 | 54.07792 | 53.57 | -0.94815 |
| 491 | 120.82 | 3.2657 | 54.07792 | 53.57 | -0.94815 |
| 492 | 120.82 | 3.2658 | 54.07792 | 53.58 | -0.92931 |
| 493 | 120.82 | 3.2657 | 54.07792 | 53.57 | -0.94815 |
| 494 | 120.84 | 3.266 | 54.12987 | 53.6 | -0.98856 |
| 495 | 120.83 | 3.2661 | 54.1039 | 53.61 | -0.92128 |
| 496 | 120.82 | 3.2662 | 54.07792 | 53.62 | -0.85401 |
| 497 | 120.83 | 3.2662 | 54.1039 | 53.62 | -0.90245 |
| 498 | 120.82 | 3.2661 | 54.07792 | 53.61 | -0.87283 |
| 499 | 120.82 | 3.2662 | 54.07792 | 53.62 | -0.85401 |
| 500 | 120.82 | 3.2661 | 54.07792 | 53.61 | -0.87283 |
| 501 | 120.83 | 3.2665 | 54.1039 | 53.65 | -0.84603 |
| 502 | 120.82 | 3.2666 | 54.07792 | 53.66 | -0.77883 |
| 503 | 120.83 | 3.267 | 54.1039 | 53.7 | -0.75213 |
| 504 | 120.83 | 3.2668 | 54.1039 | 53.68 | -0.78967 |
| 505 | 120.85 | 3.2669 | 54.15584 | 53.69 | -0.86766 |
| 506 | 120.82 | 3.2672 | 54.07792 | 53.72 | -0.66627 |
| 507 | 120.83 | 3.2672 | 54.1039 | 53.72 | -0.71462 |
| 508 | 120.85 | 3.2677 | 54.15584 | 53.77 | -0.71758 |
| 509 | 120.84 | 3.2685 | 54.12987 | 53.85 | -0.51972 |
| 510 | 120.83 | 3.269 | 54.1039 | 53.9 | -0.37829 |
| 511 | 120.84 | 3.2686 | 54.12987 | 53.86 | -0.50106 |
| 512 | 120.85 | 3.2692 | 54.15584 | 53.92 | -0.4374 |
| 513 | 120.83 | 3.2699 | 54.1039 | 53.99 | -0.21096 |
| 514 | 120.84 | 3.2698 | 54.12987 | 53.98 | -0.27764 |
| 515 | 120.82 | 3.2701 | 54.07792 | 54.01 | -0.12576 |
| 516 | 120.86 | 3.2714 | 54.18182 | 54.14 | -0.07724 |
| 517 | 120.86 | 3.2712 | 54.18182 | 54.12 | -0.11422 |
| 518 | 120.85 | 3.2713 | 54.15584 | 54.13 | -0.04774 |
| 519 | 120.85 | 3.2717 | 54.15584 | 54.17 | 0.026132 |
| 520 | 120.85 | 3.2716 | 54.15584 | 54.16 | 0.007673 |
| 521 | 120.88 | 3.2719 | 54.23377 | 54.19 | -0.08076 |
| 522 | 120.88 | 3.272 | 54.23377 | 54.2 | -0.0623 |
| 523 | 120.84 | 3.2726 | 54.12987 | 54.26 | 0.239827 |
| 524 | 120.84 | 3.2725 | 54.12987 | 54.25 | 0.221438 |
| 525 | 120.86 | 3.2732 | 54.18182 | 54.32 | 0.254385 |
| 526 | 120.86 | 3.2742 | 54.18182 | 54.42 | 0.437673 |
| 527 | 120.85 | 3.2765 | 54.15584 | 54.65 | 0.904219 |
| 528 | 120.88 | 3.2762 | 54.23377 | 54.62 | 0.707129 |
| 529 | 120.87 | 3.2757 | 54.20779 | 54.57 | 0.663749 |
| 530 | 120.86 | 3.2771 | 54.18182 | 54.71 | 0.965421 |
| 531 | 120.9 | 3.2759 | 54.28571 | 54.59 | 0.557402 |
| 532 | 120.88 | 3.2772 | 54.23377 | 54.72 | 0.888585 |
| 533 | 120.87 | 3.277 | 54.20779 | 54.7 | 0.899831 |
| 534 | 120.88 | 3.2776 | 54.23377 | 54.76 | 0.960982 |
| 535 | 120.87 | 3.278 | 54.20779 | 54.8 | 1.080671 |
| 536 | 120.88 | 3.2783 | 54.23377 | 54.83 | 1.087423 |
| 537 | 120.9 | 3.2786 | 54.28571 | 54.86 | 1.04682 |
| 538 | 120.89 | 3.277 | 54.25974 | 54.7 | 0.804862 |
| 539 | 120.9 | 3.2793 | 54.28571 | 54.93 | 1.172921 |
| 540 | 120.89 | 3.2796 | 54.25974 | 54.96 | 1.274126 |
| 541 | 120.9 | 3.2807 | 54.28571 | 55.07 | 1.424161 |
| 542 | 120.91 | 3.2812 | 54.31169 | 55.12 | 1.466458 |
| 543 | 120.92 | 3.2827 | 54.33766 | 55.27 | 1.686878 |
| 544 | 120.9 | 3.2819 | 54.28571 | 55.19 | 1.638496 |
| 545 | 120.93 | 3.2822 | 54.36364 | 55.22 | 1.550822 |
| 546 | 120.89 | 3.2804 | 54.25974 | 55.04 | 1.417623 |
| 547 | 120.91 | 3.2835 | 54.31169 | 55.35 | 1.875902 |
| 548 | 120.91 | 3.2826 | 54.31169 | 55.26 | 1.716091 |
| 549 | 120.9 | 3.2845 | 54.28571 | 55.45 | 2.099704 |
| 550 | 120.92 | 3.2856 | 54.33766 | 55.56 | 2.200032 |
| 551 | 120.89 | 3.2839 | 54.25974 | 55.39 | 2.040548 |
| 552 | 120.91 | 3.286 | 54.31169 | 55.6 | 2.317107 |
| 553 | 120.92 | 3.2855 | 54.33766 | 55.55 | 2.182426 |
| 554 | 120.92 | 3.2864 | 54.33766 | 55.64 | 2.34065 |
| 555 | 120.91 | 3.2878 | 54.31169 | 55.78 | 2.632326 |
| 556 | 120.94 | 3.2876 | 54.38961 | 55.76 | 2.457657 |
| 557 | 120.94 | 3.2882 | 54.38961 | 55.82 | 2.562504 |
| 558 | 120.95 | 3.2886 | 54.41558 | 55.86 | 2.585778 |
| 559 | 120.93 | 3.2882 | 54.36364 | 55.82 | 2.609036 |
| 560 | 120.94 | 3.2892 | 54.38961 | 55.92 | 2.736748 |
| 561 | 120.93 | 3.2892 | 54.36364 | 55.92 | 2.783197 |
| 562 | 120.93 | 3.2897 | 54.36364 | 55.97 | 2.870044 |
| 563 | 120.95 | 3.2925 | 54.41558 | 56.25 | 3.261183 |
| 564 | 120.94 | 3.2913 | 54.38961 | 56.13 | 3.100641 |
| 565 | 120.98 | 3.2932 | 54.49351 | 56.32 | 3.243064 |
| 566 | 120.93 | 3.2934 | 54.36364 | 56.34 | 3.507923 |
| 567 | 120.96 | 3.294 | 54.44156 | 56.4 | 3.472414 |
| 568 | 120.94 | 3.2942 | 54.38961 | 56.42 | 3.598705 |
| 569 | 120.94 | 3.294 | 54.38961 | 56.4 | 3.564521 |
| 570 | 120.97 | 3.2944 | 54.46753 | 56.44 | 3.494804 |
| 571 | 120.96 | 3.2958 | 54.44156 | 56.58 | 3.779501 |
| 572 | 120.96 | 3.2966 | 54.44156 | 56.66 | 3.915357 |
| 573 | 120.98 | 3.297 | 54.49351 | 56.7 | 3.891523 |
| 574 | 120.97 | 3.2974 | 54.46753 | 56.74 | 4.005054 |
| 575 | 121 | 3.298 | 54.54545 | 56.8 | 3.96927 |
| 576 | 121.01 | 3.2988 | 54.57143 | 56.88 | 4.05867 |
| 577 | 121.02 | 3.2982 | 54.5974 | 56.82 | 3.911646 |
| 578 | 121.01 | 3.2986 | 54.57143 | 56.86 | 4.024923 |
| 579 | 121.01 | 3.2987 | 54.57143 | 56.87 | 4.0418 |
| 580 | 121.02 | 3.3004 | 54.5974 | 57.04 | 4.282254 |
| 581 | 121.02 | 3.3006 | 54.5974 | 57.06 | 4.315803 |
| 582 | 121.02 | 3.301 | 54.5974 | 57.1 | 4.382833 |
| 583 | 121.02 | 3.3008 | 54.5974 | 57.08 | 4.34933 |
| 584 | 121.04 | 3.3025 | 54.64935 | 57.25 | 4.542619 |
| 585 | 121.06 | 3.302 | 54.7013 | 57.2 | 4.368359 |
| 586 | 121.05 | 3.303 | 54.67532 | 57.3 | 4.580585 |
| 587 | 121.07 | 3.3031 | 54.72727 | 57.31 | 4.506591 |
| 588 | 121.01 | 3.3024 | 54.57143 | 57.24 | 4.662074 |
| 589 | 121.03 | 3.3034 | 54.62338 | 57.34 | 4.737746 |
| 590 | 121.05 | 3.3042 | 54.67532 | 57.42 | 4.779999 |
| 591 | 121.05 | 3.3048 | 54.67532 | 57.48 | 4.879393 |
| 592 | 121.07 | 3.3045 | 54.72727 | 57.45 | 4.739299 |
| 593 | 121.1 | 3.3043 | 54.80519 | 57.43 | 4.570443 |
| 594 | 121.07 | 3.3037 | 54.72727 | 57.37 | 4.606462 |
| 595 | 121.16 | 3.3041 | 54.96104 | 57.41 | 4.265739 |
| 596 | 121.07 | 3.3045 | 54.72727 | 57.45 | 4.739299 |
| 597 | 121.12 | 3.3043 | 54.85714 | 57.43 | 4.479988 |
| 598 | 121.14 | 3.3052 | 54.90909 | 57.52 | 4.539133 |
| 599 | 121.11 | 3.306 | 54.83117 | 57.6 | 4.806999 |

| Time | Therm | IC | Thermistor |  | IC Temp | % Error |
| --- | --- | --- | --- | --- | --- | --- |
| Seconds | V | V | R2 | T2 |  |  |
| 0 | 2.4975 | 2.9969 | 5324.639 | 24.95611 | 26.69 | 6.50 |
| 1 | 2.498 | 2.9968 | 5322.509 | 24.96489 | 26.68 | 6.43 |
| 2 | 2.4987 | 2.9969 | 5319.529 | 24.97717 | 26.69 | 6.42 |
| 3 | 2.4982 | 2.997 | 5321.658 | 24.9684 | 26.7 | 6.49 |
| 4 | 2.4986 | 2.9969 | 5319.955 | 24.97542 | 26.69 | 6.42 |
| 5 | 2.499 | 2.997 | 5318.253 | 24.98244 | 26.7 | 6.43 |
| 6 | 2.4988 | 2.997 | 5319.104 | 24.97893 | 26.7 | 6.45 |
| 7 | 2.5006 | 2.9971 | 5311.45 | 25.01054 | 26.71 | 6.36 |
| 8 | 2.501 | 2.9971 | 5309.75 | 25.01756 | 26.71 | 6.34 |
| 9 | 2.5027 | 2.9972 | 5302.534 | 25.04742 | 26.72 | 6.26 |
| 10 | 2.503 | 2.9972 | 5301.262 | 25.05269 | 26.72 | 6.24 |
| 11 | 2.5048 | 2.9971 | 5293.633 | 25.08431 | 26.71 | 6.09 |
| 12 | 2.5052 | 2.9971 | 5291.94 | 25.09134 | 26.71 | 6.06 |
| 13 | 2.5058 | 2.9973 | 5289.4 | 25.10188 | 26.73 | 6.09 |
| 14 | 2.507 | 2.9972 | 5284.325 | 25.12297 | 26.72 | 5.98 |
| 15 | 2.5085 | 2.9972 | 5277.987 | 25.14933 | 26.72 | 5.88 |
| 16 | 2.5089 | 2.9971 | 5276.299 | 25.15636 | 26.71 | 5.82 |
| 17 | 2.5107 | 2.9973 | 5268.706 | 25.188 | 26.73 | 5.77 |
| 18 | 2.511 | 2.9974 | 5267.442 | 25.19328 | 26.74 | 5.78 |
| 19 | 2.5122 | 2.9975 | 5262.387 | 25.21438 | 26.75 | 5.74 |
| 20 | 2.5131 | 2.9974 | 5258.6 | 25.23021 | 26.74 | 5.65 |
| 21 | 2.5134 | 2.9976 | 5257.338 | 25.23548 | 26.76 | 5.70 |
| 22 | 2.5138 | 2.9976 | 5255.656 | 25.24252 | 26.76 | 5.67 |
| 23 | 2.5142 | 2.9975 | 5253.974 | 25.24955 | 26.75 | 5.61 |
| 24 | 2.5143 | 2.9975 | 5253.554 | 25.25131 | 26.75 | 5.60 |
| 25 | 2.5145 | 2.9976 | 5252.713 | 25.25483 | 26.76 | 5.62 |
| 26 | 2.5147 | 2.9976 | 5251.873 | 25.25835 | 26.76 | 5.61 |
| 27 | 2.5149 | 2.9976 | 5251.032 | 25.26186 | 26.76 | 5.60 |
| 28 | 2.5154 | 2.9977 | 5248.932 | 25.27066 | 26.77 | 5.60 |
| 29 | 2.5162 | 2.9977 | 5245.574 | 25.28473 | 26.77 | 5.55 |
| 30 | 2.5176 | 2.9977 | 5239.702 | 25.30937 | 26.77 | 5.46 |
| 31 | 2.521 | 2.9978 | 5225.468 | 25.36921 | 26.78 | 5.27 |
| 32 | 2.5219 | 2.9978 | 5221.707 | 25.38505 | 26.78 | 5.21 |
| 33 | 2.5246 | 2.9977 | 5210.44 | 25.4326 | 26.77 | 5.00 |
| 34 | 2.5284 | 2.9978 | 5194.622 | 25.49954 | 26.78 | 4.78 |
| 35 | 2.5347 | 2.9978 | 5168.503 | 25.61059 | 26.78 | 4.37 |
| 36 | 2.5399 | 2.9978 | 5147.042 | 25.70232 | 26.78 | 4.02 |
| 37 | 2.5429 | 2.9978 | 5134.7 | 25.75527 | 26.78 | 3.83 |
| 38 | 2.5475 | 2.9978 | 5115.833 | 25.8365 | 26.78 | 3.52 |
| 39 | 2.5514 | 2.9978 | 5099.89 | 25.90541 | 26.78 | 3.27 |
| 40 | 2.555 | 2.9978 | 5085.217 | 25.96905 | 26.78 | 3.03 |
| 41 | 2.5571 | 2.9978 | 5076.677 | 26.00619 | 26.78 | 2.89 |
| 42 | 2.558 | 2.9978 | 5073.021 | 26.02211 | 26.78 | 2.83 |
| 43 | 2.5586 | 2.9978 | 5070.585 | 26.03272 | 26.78 | 2.79 |
| 44 | 2.5593 | 2.9978 | 5067.745 | 26.04511 | 26.78 | 2.74 |
| 45 | 2.5597 | 2.9978 | 5066.123 | 26.05218 | 26.78 | 2.72 |
| 46 | 2.56 | 2.9978 | 5064.906 | 26.05749 | 26.78 | 2.70 |
| 47 | 2.5611 | 2.9978 | 5060.448 | 26.07696 | 26.78 | 2.63 |
| 48 | 2.5623 | 2.9978 | 5055.59 | 26.09819 | 26.78 | 2.55 |
| 49 | 2.5655 | 2.9978 | 5042.656 | 26.15485 | 26.78 | 2.33 |
| 50 | 2.5693 | 2.9978 | 5027.338 | 26.22215 | 26.78 | 2.08 |
| 51 | 2.5753 | 2.9978 | 5003.245 | 26.3285 | 26.78 | 1.69 |
| 52 | 2.5814 | 2.9979 | 4978.864 | 26.43671 | 26.79 | 1.32 |
| 53 | 2.5875 | 2.9979 | 4954.599 | 26.54502 | 26.79 | 0.91 |
| 54 | 2.5921 | 2.9979 | 4936.376 | 26.62676 | 26.79 | 0.61 |
| 55 | 2.5979 | 2.998 | 4913.491 | 26.7299 | 26.8 | 0.26 |
| 56 | 2.6008 | 2.998 | 4902.087 | 26.78151 | 26.8 | 0.07 |
| 57 | 2.6022 | 2.9981 | 4896.591 | 26.80643 | 26.81 | 0.01 |
| 58 | 2.6036 | 2.9981 | 4891.101 | 26.83136 | 26.81 | -0.08 |
| 59 | 2.6042 | 2.9981 | 4888.749 | 26.84204 | 26.81 | -0.12 |
| 60 | 2.6045 | 2.9983 | 4887.574 | 26.84738 | 26.83 | -0.06 |
| 61 | 2.605 | 2.9984 | 4885.616 | 26.85629 | 26.84 | -0.06 |
| 62 | 2.6058 | 2.9989 | 4882.485 | 26.87054 | 26.89 | 0.07 |
| 63 | 2.607 | 2.9992 | 4877.791 | 26.89191 | 26.92 | 0.10 |
| 64 | 2.6106 | 2.9992 | 4863.737 | 26.95606 | 26.92 | -0.13 |
| 65 | 2.6166 | 2.9991 | 4840.399 | 27.06307 | 26.91 | -0.57 |
| 66 | 2.6239 | 2.9996 | 4812.148 | 27.19339 | 26.96 | -0.87 |
| 67 | 2.6298 | 2.9999 | 4789.43 | 27.29882 | 26.99 | -1.14 |
| 68 | 2.6376 | 3.0002 | 4759.552 | 27.43837 | 27.02 | -1.55 |
| 69 | 2.6442 | 3.0003 | 4734.408 | 27.55659 | 27.03 | -1.95 |
| 70 | 2.6473 | 3.0014 | 4722.641 | 27.61216 | 27.14 | -1.74 |
| 71 | 2.6484 | 3.0028 | 4718.472 | 27.63188 | 27.28 | -1.29 |
| 72 | 2.649 | 3.0029 | 4716.2 | 27.64264 | 27.29 | -1.29 |
| 73 | 2.6494 | 3.0035 | 4714.686 | 27.64982 | 27.35 | -1.10 |
| 74 | 2.6495 | 3.0051 | 4714.307 | 27.65161 | 27.51 | -0.51 |
| 75 | 2.6495 | 3.0057 | 4714.307 | 27.65161 | 27.57 | -0.30 |
| 76 | 2.6495 | 3.0076 | 4714.307 | 27.65161 | 27.76 | 0.39 |
| 77 | 2.6496 | 3.0092 | 4713.929 | 27.6534 | 27.92 | 0.95 |
| 78 | 2.6498 | 3.0098 | 4713.172 | 27.65699 | 27.98 | 1.15 |
| 79 | 2.6506 | 3.0116 | 4710.145 | 27.67134 | 28.16 | 1.74 |
| 80 | 2.6529 | 3.0148 | 4701.455 | 27.71261 | 28.48 | 2.69 |
| 81 | 2.6549 | 3.0153 | 4693.91 | 27.74851 | 28.53 | 2.74 |
| 82 | 2.6586 | 3.0172 | 4679.982 | 27.81496 | 28.72 | 3.15 |
| 83 | 2.663 | 3.0167 | 4663.469 | 27.89404 | 28.67 | 2.71 |
| 84 | 2.6702 | 3.0185 | 4636.566 | 28.02356 | 28.85 | 2.86 |
| 85 | 2.6776 | 3.0212 | 4609.065 | 28.15685 | 29.12 | 3.31 |
| 86 | 2.6834 | 3.0243 | 4587.617 | 28.26144 | 29.43 | 3.97 |
| 87 | 2.6877 | 3.0275 | 4571.776 | 28.33906 | 29.75 | 4.74 |
| 88 | 2.6909 | 3.0292 | 4560.02 | 28.39685 | 29.92 | 5.09 |
| 89 | 2.6932 | 3.0298 | 4551.587 | 28.43842 | 29.98 | 5.14 |
| 90 | 2.6942 | 3.0312 | 4547.926 | 28.45649 | 30.12 | 5.52 |
| 91 | 2.6945 | 3.0328 | 4546.828 | 28.46192 | 30.28 | 6.00 |
| 92 | 2.6948 | 3.0322 | 4545.73 | 28.46734 | 30.22 | 5.80 |
| 93 | 2.6953 | 3.0348 | 4543.901 | 28.47638 | 30.48 | 6.57 |
| 94 | 2.6964 | 3.0348 | 4539.879 | 28.49627 | 30.48 | 6.51 |
| 95 | 2.701 | 3.0358 | 4523.097 | 28.5795 | 30.58 | 6.54 |
| 96 | 2.7079 | 3.0362 | 4498.031 | 28.70447 | 30.62 | 6.26 |
| 97 | 2.7142 | 3.038 | 4475.257 | 28.81872 | 30.8 | 6.43 |
| 98 | 2.7199 | 3.0394 | 4454.741 | 28.9222 | 30.94 | 6.52 |
| 99 | 2.727 | 3.0398 | 4429.308 | 29.05126 | 30.98 | 6.23 |
| 100 | 2.7341 | 3.04 | 4404.006 | 29.18049 | 31 | 5.87 |
| 101 | 2.7375 | 3.0408 | 4391.936 | 29.24244 | 31.08 | 5.91 |
| 102 | 2.7413 | 3.0419 | 4378.482 | 29.31173 | 31.19 | 6.02 |
| 103 | 2.7451 | 3.0427 | 4365.065 | 29.38107 | 31.27 | 6.04 |
| 104 | 2.7503 | 3.043 | 4346.764 | 29.47604 | 31.3 | 5.83 |
| 105 | 2.7568 | 3.0433 | 4323.986 | 29.5949 | 31.33 | 5.54 |
| 106 | 2.7617 | 3.0438 | 4306.886 | 29.6846 | 31.38 | 5.40 |
| 107 | 2.7659 | 3.0448 | 4292.276 | 29.76156 | 31.48 | 5.46 |
| 108 | 2.7724 | 3.0459 | 4269.754 | 29.8808 | 31.59 | 5.41 |
| 109 | 2.7768 | 3.0455 | 4254.568 | 29.96161 | 31.55 | 5.03 |
| 110 | 2.7808 | 3.0471 | 4240.804 | 30.03514 | 31.71 | 5.28 |
| 111 | 2.7828 | 3.0483 | 4233.937 | 30.07192 | 31.83 | 5.52 |
| 112 | 2.7835 | 3.05 | 4231.536 | 30.0848 | 32 | 5.98 |
| 113 | 2.7845 | 3.0522 | 4228.108 | 30.1032 | 32.22 | 6.57 |
| 114 | 2.7852 | 3.0548 | 4225.71 | 30.11609 | 32.48 | 7.28 |
| 115 | 2.7865 | 3.059 | 4221.259 | 30.14002 | 32.9 | 8.39 |
| 116 | 2.789 | 3.0594 | 4212.712 | 30.18606 | 32.94 | 8.36 |
| 117 | 2.7946 | 3.0613 | 4193.622 | 30.28928 | 33.13 | 8.57 |
| 118 | 2.8001 | 3.0609 | 4174.947 | 30.39078 | 33.09 | 8.16 |
| 119 | 2.8079 | 3.0652 | 4148.588 | 30.53494 | 33.52 | 8.91 |
| 120 | 2.8168 | 3.0633 | 4118.69 | 30.69974 | 33.33 | 7.89 |
| 121 | 2.8234 | 3.0643 | 4096.64 | 30.82216 | 33.43 | 7.80 |
| 122 | 2.8263 | 3.0688 | 4086.984 | 30.876 | 33.88 | 8.87 |
| 123 | 2.8277 | 3.0705 | 4082.329 | 30.90201 | 34.05 | 9.25 |
| 124 | 2.8289 | 3.0714 | 4078.343 | 30.92431 | 34.14 | 9.42 |
| 125 | 2.8294 | 3.0715 | 4076.684 | 30.93361 | 34.15 | 9.42 |
| 126 | 2.8295 | 3.0736 | 4076.352 | 30.93546 | 34.36 | 9.97 |
| 127 | 2.8295 | 3.076 | 4076.352 | 30.93546 | 34.6 | 10.59 |
| 128 | 2.8295 | 3.0779 | 4076.352 | 30.93546 | 34.79 | 11.08 |
| 129 | 2.8295 | 3.0794 | 4076.352 | 30.93546 | 34.94 | 11.46 |
| 130 | 2.8297 | 3.0807 | 4075.688 | 30.93918 | 35.07 | 11.78 |
| 131 | 2.8302 | 3.0809 | 4074.029 | 30.94848 | 35.09 | 11.80 |
| 132 | 2.8317 | 3.0824 | 4069.056 | 30.97637 | 35.24 | 12.10 |
| 133 | 2.8329 | 3.0838 | 4065.082 | 30.99869 | 35.38 | 12.38 |
| 134 | 2.8375 | 3.0838 | 4049.877 | 31.08431 | 35.38 | 12.14 |
| 135 | 2.8428 | 3.0845 | 4032.419 | 31.18307 | 35.45 | 12.04 |
| 136 | 2.849 | 3.0849 | 4012.079 | 31.29876 | 35.49 | 11.81 |
| 137 | 2.8541 | 3.0854 | 3995.415 | 31.39404 | 35.54 | 11.67 |
| 138 | 2.86 | 3.0855 | 3976.21 | 31.50443 | 35.55 | 11.38 |
| 139 | 2.8665 | 3.0858 | 3955.144 | 31.62621 | 35.58 | 11.11 |
| 140 | 2.8703 | 3.0862 | 3942.872 | 31.6975 | 35.62 | 11.01 |
| 141 | 2.872 | 3.0863 | 3937.393 | 31.72941 | 35.63 | 10.95 |
| 142 | 2.8736 | 3.0865 | 3932.242 | 31.75946 | 35.65 | 10.91 |
| 143 | 2.8743 | 3.0865 | 3929.99 | 31.77261 | 35.65 | 10.88 |
| 144 | 2.8747 | 3.0866 | 3928.704 | 31.78013 | 35.66 | 10.88 |
| 145 | 2.8751 | 3.0867 | 3927.418 | 31.78764 | 35.67 | 10.88 |
| 146 | 2.8764 | 3.0868 | 3923.241 | 31.81208 | 35.68 | 10.84 |
| 147 | 2.8789 | 3.0868 | 3915.219 | 31.85908 | 35.68 | 10.71 |
| 148 | 2.882 | 3.0869 | 3905.292 | 31.91741 | 35.69 | 10.57 |
| 149 | 2.89 | 3.0869 | 3879.772 | 32.06814 | 35.69 | 10.15 |
| 150 | 2.8978 | 3.087 | 3855.025 | 32.2154 | 35.7 | 9.76 |
| 151 | 2.9032 | 3.087 | 3837.97 | 32.31752 | 35.7 | 9.47 |
| 152 | 2.9087 | 3.0872 | 3820.665 | 32.42168 | 35.72 | 9.23 |
| 153 | 2.913 | 3.0873 | 3807.181 | 32.50321 | 35.73 | 9.03 |
| 154 | 2.9167 | 3.0873 | 3795.61 | 32.57344 | 35.73 | 8.83 |
| 155 | 2.918 | 3.0874 | 3791.552 | 32.59813 | 35.74 | 8.79 |
| 156 | 2.919 | 3.0876 | 3788.432 | 32.61713 | 35.76 | 8.79 |
| 157 | 2.9196 | 3.0878 | 3786.562 | 32.62853 | 35.78 | 8.81 |
| 158 | 2.9199 | 3.0879 | 3785.627 | 32.63423 | 35.79 | 8.82 |
| 159 | 2.9206 | 3.0881 | 3783.446 | 32.64754 | 35.81 | 8.83 |
| 160 | 2.9224 | 3.0884 | 3777.842 | 32.68176 | 35.84 | 8.81 |
| 161 | 2.9251 | 3.0893 | 3769.45 | 32.73313 | 35.93 | 8.90 |
| 162 | 2.9313 | 3.0894 | 3750.238 | 32.85123 | 35.94 | 8.59 |
| 163 | 2.9369 | 3.0896 | 3732.954 | 32.95807 | 35.96 | 8.35 |
| 164 | 2.9431 | 3.0904 | 3713.896 | 33.07654 | 36.04 | 8.22 |
| 165 | 2.9481 | 3.0909 | 3698.584 | 33.17222 | 36.09 | 8.08 |
| 166 | 2.955 | 3.0914 | 3677.54 | 33.30448 | 36.14 | 7.85 |
| 167 | 2.9591 | 3.0938 | 3665.081 | 33.38319 | 36.38 | 8.24 |
| 168 | 2.9619 | 3.0939 | 3656.593 | 33.437 | 36.39 | 8.11 |
| 169 | 2.9632 | 3.0942 | 3652.658 | 33.46199 | 36.42 | 8.12 |
| 170 | 2.964 | 3.0954 | 3650.238 | 33.47738 | 36.54 | 8.38 |
| 171 | 2.9643 | 3.0963 | 3649.33 | 33.48315 | 36.63 | 8.59 |
| 172 | 2.9647 | 3.0988 | 3648.121 | 33.49084 | 36.88 | 9.19 |
| 173 | 2.9654 | 3.1006 | 3646.005 | 33.50431 | 37.06 | 9.59 |
| 174 | 2.966 | 3.1003 | 3644.193 | 33.51586 | 37.03 | 9.49 |
| 175 | 2.9678 | 3.1025 | 3638.76 | 33.55051 | 37.25 | 9.93 |
| 176 | 2.9712 | 3.1041 | 3628.515 | 33.61601 | 37.41 | 10.14 |
| 177 | 2.977 | 3.1059 | 3611.092 | 33.72788 | 37.59 | 10.27 |
| 178 | 2.9845 | 3.1086 | 3588.664 | 33.87283 | 37.86 | 10.53 |
| 179 | 2.9895 | 3.1095 | 3573.774 | 33.96963 | 37.95 | 10.49 |
| 180 | 2.9972 | 3.1111 | 3550.941 | 34.11898 | 38.11 | 10.47 |
| 181 | 3.0012 | 3.1134 | 3539.125 | 34.19669 | 38.34 | 10.81 |
| 182 | 3.0054 | 3.114 | 3526.753 | 34.27839 | 38.4 | 10.73 |
| 183 | 3.0078 | 3.1149 | 3519.699 | 34.32513 | 38.49 | 10.82 |
| 184 | 3.0087 | 3.1179 | 3517.057 | 34.34266 | 38.79 | 11.47 |
| 185 | 3.0093 | 3.1186 | 3515.296 | 34.35435 | 38.86 | 11.59 |
| 186 | 3.0094 | 3.118 | 3515.002 | 34.3563 | 38.8 | 11.45 |
| 187 | 3.0095 | 3.1186 | 3514.709 | 34.35825 | 38.86 | 11.58 |
| 188 | 3.0096 | 3.1211 | 3514.416 | 34.3602 | 39.11 | 12.14 |
| 189 | 3.0096 | 3.123 | 3514.416 | 34.3602 | 39.3 | 12.57 |
| 190 | 3.0096 | 3.1242 | 3514.416 | 34.3602 | 39.42 | 12.84 |
| 191 | 3.0096 | 3.125 | 3514.416 | 34.3602 | 39.5 | 13.01 |
| 192 | 3.0098 | 3.1256 | 3513.829 | 34.36409 | 39.56 | 13.13 |
| 193 | 3.0101 | 3.1267 | 3512.949 | 34.36994 | 39.67 | 13.36 |
| 194 | 3.0108 | 3.1268 | 3510.897 | 34.38359 | 39.68 | 13.35 |
| 195 | 3.012 | 3.1273 | 3507.381 | 34.40699 | 39.73 | 13.40 |
| 196 | 3.0135 | 3.1275 | 3502.99 | 34.43625 | 39.75 | 13.37 |
| 197 | 3.0166 | 3.1282 | 3493.929 | 34.49676 | 39.82 | 13.37 |
| 198 | 3.0197 | 3.129 | 3484.887 | 34.55733 | 39.9 | 13.39 |
| 199 | 3.0242 | 3.1294 | 3471.795 | 34.64535 | 39.94 | 13.26 |
| 200 | 3.0292 | 3.1298 | 3457.293 | 34.7433 | 39.98 | 13.10 |
| 201 | 3.0341 | 3.1305 | 3443.127 | 34.83943 | 40.05 | 13.01 |
| 202 | 3.0393 | 3.1305 | 3428.145 | 34.94161 | 40.05 | 12.76 |
| 203 | 3.0456 | 3.1304 | 3410.061 | 35.06562 | 40.04 | 12.42 |
| 204 | 3.0484 | 3.1306 | 3402.048 | 35.12081 | 40.06 | 12.33 |
| 205 | 3.0506 | 3.1311 | 3395.762 | 35.16421 | 40.11 | 12.33 |
| 206 | 3.0525 | 3.1312 | 3390.341 | 35.20171 | 40.12 | 12.26 |
| 207 | 3.0536 | 3.1314 | 3387.205 | 35.22344 | 40.14 | 12.25 |
| 208 | 3.0543 | 3.1318 | 3385.211 | 35.23727 | 40.18 | 12.30 |
| 209 | 3.0544 | 3.1321 | 3384.926 | 35.23924 | 40.21 | 12.36 |
| 210 | 3.0546 | 3.1322 | 3384.357 | 35.24319 | 40.22 | 12.37 |
| 211 | 3.0549 | 3.1326 | 3383.502 | 35.24912 | 40.26 | 12.45 |
| 212 | 3.0556 | 3.1328 | 3381.51 | 35.26296 | 40.28 | 12.46 |
| 213 | 3.0569 | 3.133 | 3377.812 | 35.28866 | 40.3 | 12.44 |
| 214 | 3.0605 | 3.1333 | 3367.588 | 35.35988 | 40.33 | 12.32 |
| 215 | 3.063 | 3.1335 | 3360.502 | 35.40939 | 40.35 | 12.24 |
| 216 | 3.0702 | 3.1342 | 3340.159 | 35.5522 | 40.42 | 12.04 |
| 217 | 3.0769 | 3.1349 | 3321.315 | 35.68538 | 40.49 | 11.87 |
| 218 | 3.0822 | 3.1363 | 3306.466 | 35.79094 | 40.63 | 11.91 |
| 219 | 3.0856 | 3.1364 | 3296.967 | 35.85876 | 40.64 | 11.76 |
| 220 | 3.0907 | 3.1359 | 3282.758 | 35.96062 | 40.59 | 11.41 |
| 221 | 3.0934 | 3.137 | 3275.255 | 36.01462 | 40.7 | 11.51 |
| 222 | 3.0969 | 3.14 | 3265.547 | 36.08468 | 41 | 11.99 |
| 223 | 3.098 | 3.1398 | 3262.501 | 36.10672 | 40.98 | 11.89 |
| 224 | 3.0988 | 3.1409 | 3260.287 | 36.12276 | 41.09 | 12.09 |
| 225 | 3.0994 | 3.1436 | 3258.627 | 36.13478 | 41.36 | 12.63 |
| 226 | 3.1 | 3.1434 | 3256.968 | 36.14681 | 41.34 | 12.56 |
| 227 | 3.1005 | 3.1455 | 3255.586 | 36.15684 | 41.55 | 12.98 |
| 228 | 3.1014 | 3.1484 | 3253.099 | 36.17489 | 41.84 | 13.54 |
| 229 | 3.1032 | 3.1481 | 3248.129 | 36.21101 | 41.81 | 13.39 |
| 230 | 3.1051 | 3.1494 | 3242.89 | 36.24916 | 41.94 | 13.57 |
| 231 | 3.1084 | 3.1508 | 3233.806 | 36.31548 | 42.08 | 13.70 |
| 232 | 3.1141 | 3.1518 | 3218.16 | 36.43021 | 42.18 | 13.63 |
| 233 | 3.119 | 3.1525 | 3204.756 | 36.52901 | 42.25 | 13.54 |
| 234 | 3.1248 | 3.1541 | 3188.944 | 36.64618 | 42.41 | 13.59 |
| 235 | 3.1292 | 3.1553 | 3176.988 | 36.73522 | 42.53 | 13.63 |
| 236 | 3.1318 | 3.1602 | 3169.939 | 36.78789 | 43.02 | 14.49 |
| 237 | 3.1366 | 3.1606 | 3156.956 | 36.88527 | 43.06 | 14.34 |
| 238 | 3.1391 | 3.1611 | 3150.209 | 36.93605 | 43.11 | 14.32 |
| 239 | 3.1411 | 3.1623 | 3144.82 | 36.97671 | 43.23 | 14.47 |
| 240 | 3.1425 | 3.1644 | 3141.052 | 37.00518 | 43.44 | 14.81 |
| 241 | 3.1433 | 3.1652 | 3138.9 | 37.02146 | 43.52 | 14.93 |
| 242 | 3.1442 | 3.1661 | 3136.48 | 37.03978 | 43.61 | 15.07 |
| 243 | 3.1447 | 3.1672 | 3135.137 | 37.04996 | 43.72 | 15.26 |
| 244 | 3.1449 | 3.1698 | 3134.599 | 37.05403 | 43.98 | 15.75 |
| 245 | 3.1455 | 3.17 | 3132.988 | 37.06625 | 44 | 15.76 |
| 246 | 3.1464 | 3.1699 | 3130.572 | 37.08459 | 43.99 | 15.70 |
| 247 | 3.1487 | 3.1714 | 3124.403 | 37.13146 | 44.14 | 15.88 |
| 248 | 3.1517 | 3.1714 | 3116.371 | 37.19267 | 44.14 | 15.74 |
| 249 | 3.1543 | 3.1727 | 3109.422 | 37.24576 | 44.27 | 15.87 |
| 250 | 3.1574 | 3.1728 | 3101.152 | 37.30913 | 44.28 | 15.74 |
| 251 | 3.1609 | 3.1736 | 3091.834 | 37.38076 | 44.36 | 15.73 |
| 252 | 3.1669 | 3.1739 | 3075.908 | 37.50377 | 44.39 | 15.51 |
| 253 | 3.1725 | 3.1744 | 3061.099 | 37.61882 | 44.44 | 15.35 |
| 254 | 3.1756 | 3.1748 | 3052.923 | 37.68261 | 44.48 | 15.28 |
| 255 | 3.1788 | 3.1753 | 3044.5 | 37.74853 | 44.53 | 15.23 |
| 256 | 3.1816 | 3.1753 | 3037.144 | 37.80627 | 44.53 | 15.10 |
| 257 | 3.1847 | 3.1755 | 3029.015 | 37.87027 | 44.55 | 14.99 |
| 258 | 3.1859 | 3.1755 | 3025.873 | 37.89507 | 44.55 | 14.94 |
| 259 | 3.1874 | 3.1758 | 3021.948 | 37.92608 | 44.58 | 14.93 |
| 260 | 3.1883 | 3.1758 | 3019.595 | 37.94469 | 44.58 | 14.88 |
| 261 | 3.1888 | 3.176 | 3018.288 | 37.95503 | 44.6 | 14.90 |
| 262 | 3.1891 | 3.176 | 3017.504 | 37.96124 | 44.6 | 14.89 |
| 263 | 3.1895 | 3.1761 | 3016.459 | 37.96952 | 44.61 | 14.89 |
| 264 | 3.1897 | 3.1761 | 3015.937 | 37.97366 | 44.61 | 14.88 |
| 265 | 3.1899 | 3.1762 | 3015.415 | 37.97779 | 44.62 | 14.89 |
| 266 | 3.1901 | 3.1762 | 3014.893 | 37.98193 | 44.62 | 14.88 |
| 267 | 3.1904 | 3.1762 | 3014.109 | 37.98814 | 44.62 | 14.86 |
| 268 | 3.1912 | 3.1762 | 3012.022 | 38.00471 | 44.62 | 14.83 |
| 269 | 3.192 | 3.1762 | 3009.935 | 38.02128 | 44.62 | 14.79 |
| 270 | 3.1925 | 3.1762 | 3008.631 | 38.03163 | 44.62 | 14.77 |
| 271 | 3.1933 | 3.1762 | 3006.546 | 38.04821 | 44.62 | 14.73 |
| 272 | 3.1946 | 3.1762 | 3003.16 | 38.07516 | 44.62 | 14.67 |
| 273 | 3.1976 | 3.1762 | 2995.357 | 38.13739 | 44.62 | 14.53 |
| 274 | 3.2003 | 3.1762 | 2988.347 | 38.19346 | 44.62 | 14.40 |
| 275 | 3.2023 | 3.1762 | 2983.161 | 38.23504 | 44.62 | 14.31 |
| 276 | 3.2043 | 3.1762 | 2977.983 | 38.27664 | 44.62 | 14.22 |
| 277 | 3.2086 | 3.1762 | 2966.87 | 38.36619 | 44.62 | 14.02 |
| 278 | 3.2111 | 3.1762 | 2960.423 | 38.41833 | 44.62 | 13.90 |
| 279 | 3.2134 | 3.1763 | 2954.501 | 38.46633 | 44.63 | 13.81 |
| 280 | 3.2164 | 3.1762 | 2946.788 | 38.52901 | 44.62 | 13.65 |
| 281 | 3.2197 | 3.1762 | 2938.322 | 38.59805 | 44.62 | 13.50 |
| 282 | 3.2229 | 3.1763 | 2930.128 | 38.66507 | 44.63 | 13.37 |
| 283 | 3.2262 | 3.1762 | 2921.695 | 38.73428 | 44.62 | 13.19 |
| 284 | 3.2283 | 3.1763 | 2916.338 | 38.77836 | 44.63 | 13.11 |
| 285 | 3.2301 | 3.1764 | 2911.752 | 38.81618 | 44.64 | 13.05 |
| 286 | 3.2323 | 3.1763 | 2906.153 | 38.86244 | 44.63 | 12.92 |
| 287 | 3.2344 | 3.1762 | 2900.816 | 38.90663 | 44.62 | 12.80 |
| 288 | 3.2352 | 3.1764 | 2898.784 | 38.92347 | 44.64 | 12.81 |
| 289 | 3.236 | 3.1764 | 2896.754 | 38.94032 | 44.64 | 12.77 |
| 290 | 3.2371 | 3.1764 | 2893.964 | 38.9635 | 44.64 | 12.72 |
| 291 | 3.2392 | 3.1765 | 2888.643 | 39.00777 | 44.65 | 12.64 |
| 292 | 3.242 | 3.1766 | 2881.558 | 39.06686 | 44.66 | 12.52 |
| 293 | 3.245 | 3.1767 | 2873.982 | 39.13025 | 44.67 | 12.40 |
| 294 | 3.2494 | 3.1767 | 2862.894 | 39.22334 | 44.67 | 12.19 |
| 295 | 3.2539 | 3.1768 | 2851.586 | 39.31873 | 44.68 | 12.00 |
| 296 | 3.2576 | 3.1771 | 2842.311 | 39.39728 | 44.71 | 11.88 |
| 297 | 3.26 | 3.1771 | 2836.307 | 39.4483 | 44.71 | 11.77 |
| 298 | 3.2645 | 3.1772 | 2825.072 | 39.54409 | 44.72 | 11.57 |
| 299 | 3.2691 | 3.1774 | 2813.619 | 39.6422 | 44.74 | 11.39 |
| 300 | 3.2712 | 3.1776 | 2808.402 | 39.68704 | 44.76 | 11.33 |
| 301 | 3.2739 | 3.1782 | 2801.703 | 39.74476 | 44.82 | 11.32 |
| 302 | 3.2757 | 3.178 | 2797.243 | 39.78327 | 44.8 | 11.20 |
| 303 | 3.2774 | 3.1784 | 2793.036 | 39.81967 | 44.84 | 11.20 |
| 304 | 3.2786 | 3.1783 | 2790.069 | 39.84538 | 44.83 | 11.12 |
| 305 | 3.2789 | 3.1791 | 2789.327 | 39.85181 | 44.91 | 11.26 |
| 306 | 3.2792 | 3.179 | 2788.586 | 39.85824 | 44.9 | 11.23 |
| 307 | 3.2794 | 3.1792 | 2788.092 | 39.86253 | 44.92 | 11.26 |
| 308 | 3.2797 | 3.1802 | 2787.351 | 39.86896 | 45.02 | 11.44 |
| 309 | 3.28 | 3.1806 | 2786.61 | 39.87539 | 45.06 | 11.51 |
| 310 | 3.2805 | 3.1804 | 2785.375 | 39.88612 | 45.04 | 11.44 |
| 311 | 3.281 | 3.1812 | 2784.141 | 39.89684 | 45.12 | 11.58 |
| 312 | 3.2819 | 3.183 | 2781.92 | 39.91615 | 45.3 | 11.88 |
| 313 | 3.2826 | 3.1847 | 2780.194 | 39.93117 | 45.47 | 12.18 |
| 314 | 3.2841 | 3.1851 | 2776.497 | 39.96338 | 45.51 | 12.19 |
| 315 | 3.2862 | 3.1862 | 2771.327 | 40.00851 | 45.62 | 12.30 |
| 316 | 3.2881 | 3.1866 | 2766.654 | 40.04937 | 45.66 | 12.29 |
| 317 | 3.2911 | 3.1875 | 2759.289 | 40.11395 | 45.75 | 12.32 |
| 318 | 3.2953 | 3.1877 | 2748.999 | 40.20451 | 45.77 | 12.16 |
| 319 | 3.2989 | 3.1886 | 2740.2 | 40.28225 | 45.86 | 12.16 |
| 320 | 3.3027 | 3.1909 | 2730.933 | 40.36445 | 46.09 | 12.42 |
| 321 | 3.3051 | 3.1911 | 2725.091 | 40.41643 | 46.11 | 12.35 |
| 322 | 3.3078 | 3.1928 | 2718.529 | 40.47497 | 46.28 | 12.54 |
| 323 | 3.311 | 3.1953 | 2710.766 | 40.54444 | 46.53 | 12.86 |
| 324 | 3.313 | 3.1967 | 2705.922 | 40.5879 | 46.67 | 13.03 |
| 325 | 3.3163 | 3.1971 | 2697.941 | 40.6597 | 46.71 | 12.95 |
| 326 | 3.3182 | 3.1971 | 2693.353 | 40.70109 | 46.71 | 12.86 |
| 327 | 3.32 | 3.1984 | 2689.012 | 40.74033 | 46.84 | 13.02 |
| 328 | 3.3209 | 3.2009 | 2686.843 | 40.75996 | 47.09 | 13.44 |
| 329 | 3.3219 | 3.2011 | 2684.435 | 40.78178 | 47.11 | 13.43 |
| 330 | 3.323 | 3.2037 | 2681.787 | 40.8058 | 47.37 | 13.86 |
| 331 | 3.3236 | 3.2053 | 2680.343 | 40.8189 | 47.53 | 14.12 |
| 332 | 3.3239 | 3.2047 | 2679.622 | 40.82546 | 47.47 | 14.00 |
| 333 | 3.3244 | 3.2056 | 2678.42 | 40.83638 | 47.56 | 14.14 |
| 334 | 3.3245 | 3.2067 | 2678.179 | 40.83856 | 47.67 | 14.33 |
| 335 | 3.3246 | 3.207 | 2677.939 | 40.84075 | 47.7 | 14.38 |
| 336 | 3.3247 | 3.208 | 2677.698 | 40.84294 | 47.8 | 14.55 |
| 337 | 3.3252 | 3.21 | 2676.497 | 40.85386 | 48 | 14.89 |
| 338 | 3.3253 | 3.2106 | 2676.257 | 40.85605 | 48.06 | 14.99 |
| 339 | 3.3258 | 3.2125 | 2675.055 | 40.86698 | 48.25 | 15.30 |
| 340 | 3.3263 | 3.2128 | 2673.854 | 40.87791 | 48.28 | 15.33 |
| 341 | 3.3273 | 3.2129 | 2671.454 | 40.89979 | 48.29 | 15.30 |
| 342 | 3.3289 | 3.2142 | 2667.616 | 40.9348 | 48.42 | 15.46 |
| 343 | 3.3313 | 3.2138 | 2661.865 | 40.98737 | 48.38 | 15.28 |
| 344 | 3.3321 | 3.2142 | 2659.95 | 41.00491 | 48.42 | 15.31 |
| 345 | 3.3348 | 3.2153 | 2653.494 | 41.06414 | 48.53 | 15.38 |
| 346 | 3.3396 | 3.2155 | 2642.043 | 41.16961 | 48.55 | 15.20 |
| 347 | 3.3445 | 3.2162 | 2630.386 | 41.27751 | 48.62 | 15.10 |
| 348 | 3.3467 | 3.2168 | 2625.164 | 41.32603 | 48.68 | 15.11 |
| 349 | 3.3494 | 3.2174 | 2618.764 | 41.38564 | 48.74 | 15.09 |
| 350 | 3.352 | 3.2177 | 2612.611 | 41.44312 | 48.77 | 15.02 |
| 351 | 3.3562 | 3.2178 | 2602.691 | 41.5361 | 48.78 | 14.85 |
| 352 | 3.3604 | 3.2184 | 2592.797 | 41.62926 | 48.84 | 14.76 |
| 353 | 3.361 | 3.2192 | 2591.385 | 41.64259 | 48.92 | 14.88 |
| 354 | 3.363 | 3.2196 | 2586.684 | 41.68702 | 48.96 | 14.85 |
| 355 | 3.364 | 3.2194 | 2584.335 | 41.70926 | 48.94 | 14.77 |
| 356 | 3.3655 | 3.2198 | 2580.815 | 41.74263 | 48.98 | 14.78 |
| 357 | 3.3667 | 3.2199 | 2578.001 | 41.76934 | 48.99 | 14.74 |
| 358 | 3.3678 | 3.2202 | 2575.423 | 41.79384 | 49.02 | 14.74 |
| 359 | 3.3684 | 3.2203 | 2574.018 | 41.8072 | 49.03 | 14.73 |
| 360 | 3.369 | 3.2207 | 2572.613 | 41.82058 | 49.07 | 14.77 |
| 361 | 3.369 | 3.2209 | 2572.613 | 41.82058 | 49.09 | 14.81 |
| 362 | 3.3693 | 3.2216 | 2571.911 | 41.82726 | 49.16 | 14.92 |
| 363 | 3.3693 | 3.2215 | 2571.911 | 41.82726 | 49.15 | 14.90 |
| 364 | 3.3694 | 3.2216 | 2571.677 | 41.82949 | 49.16 | 14.91 |
| 365 | 3.3695 | 3.2221 | 2571.443 | 41.83172 | 49.21 | 14.99 |
| 366 | 3.3695 | 3.222 | 2571.443 | 41.83172 | 49.2 | 14.98 |
| 367 | 3.3695 | 3.2224 | 2571.443 | 41.83172 | 49.24 | 15.05 |
| 368 | 3.3695 | 3.2227 | 2571.443 | 41.83172 | 49.27 | 15.10 |
| 369 | 3.3695 | 3.223 | 2571.443 | 41.83172 | 49.3 | 15.15 |
| 370 | 3.3695 | 3.2238 | 2571.443 | 41.83172 | 49.38 | 15.29 |
| 371 | 3.3695 | 3.2234 | 2571.443 | 41.83172 | 49.34 | 15.22 |
| 372 | 3.3696 | 3.224 | 2571.209 | 41.83395 | 49.4 | 15.32 |
| 373 | 3.3698 | 3.2234 | 2570.741 | 41.83841 | 49.34 | 15.20 |
| 374 | 3.3698 | 3.2242 | 2570.741 | 41.83841 | 49.42 | 15.34 |
| 375 | 3.37 | 3.2245 | 2570.273 | 41.84287 | 49.45 | 15.38 |
| 376 | 3.3702 | 3.2263 | 2569.805 | 41.84733 | 49.63 | 15.68 |
| 377 | 3.3705 | 3.226 | 2569.103 | 41.85402 | 49.6 | 15.62 |
| 378 | 3.3708 | 3.2262 | 2568.402 | 41.86071 | 49.62 | 15.64 |
| 379 | 3.3713 | 3.2262 | 2567.233 | 41.87187 | 49.62 | 15.61 |
| 380 | 3.3723 | 3.2281 | 2564.896 | 41.89419 | 49.81 | 15.89 |
| 381 | 3.3733 | 3.2286 | 2562.56 | 41.91652 | 49.86 | 15.93 |
| 382 | 3.3744 | 3.2295 | 2559.992 | 41.94109 | 49.95 | 16.03 |
| 383 | 3.3756 | 3.2293 | 2557.193 | 41.96791 | 49.93 | 15.95 |
| 384 | 3.377 | 3.2308 | 2553.93 | 41.99922 | 50.08 | 16.14 |
| 385 | 3.3777 | 3.23 | 2552.3 | 42.01488 | 50 | 15.97 |
| 386 | 3.3816 | 3.2299 | 2543.227 | 42.10224 | 49.99 | 15.78 |
| 387 | 3.3826 | 3.233 | 2540.905 | 42.12466 | 50.3 | 16.25 |
| 388 | 3.3857 | 3.2352 | 2533.712 | 42.19425 | 50.52 | 16.48 |
| 389 | 3.3902 | 3.2356 | 2523.296 | 42.29543 | 50.56 | 16.35 |
| 390 | 3.3912 | 3.2361 | 2520.985 | 42.31795 | 50.61 | 16.38 |
| 391 | 3.3912 | 3.2366 | 2520.985 | 42.31795 | 50.66 | 16.47 |
| 392 | 3.3952 | 3.2376 | 2511.754 | 42.40811 | 50.76 | 16.45 |
| 393 | 3.3975 | 3.2386 | 2506.456 | 42.46003 | 50.86 | 16.52 |
| 394 | 3.4001 | 3.2387 | 2500.476 | 42.5188 | 50.87 | 16.42 |
| 395 | 3.4023 | 3.2412 | 2495.423 | 42.56858 | 51.12 | 16.73 |
| 396 | 3.4029 | 3.2451 | 2494.046 | 42.58216 | 51.51 | 17.33 |
| 397 | 3.4062 | 3.2446 | 2486.481 | 42.65695 | 51.46 | 17.11 |
| 398 | 3.4077 | 3.2447 | 2483.048 | 42.69098 | 51.47 | 17.06 |
| 399 | 3.4101 | 3.2434 | 2477.56 | 42.74548 | 51.34 | 16.74 |
| 400 | 3.4107 | 3.247 | 2476.19 | 42.75912 | 51.7 | 17.29 |
| 401 | 3.4114 | 3.2494 | 2474.591 | 42.77503 | 51.94 | 17.65 |
| 402 | 3.4124 | 3.2477 | 2472.309 | 42.79778 | 51.77 | 17.33 |
| 403 | 3.4129 | 3.2491 | 2471.168 | 42.80915 | 51.91 | 17.53 |
| 404 | 3.4137 | 3.2476 | 2469.344 | 42.82736 | 51.76 | 17.26 |
| 405 | 3.4138 | 3.2494 | 2469.116 | 42.82963 | 51.94 | 17.54 |
| 406 | 3.4141 | 3.2518 | 2468.432 | 42.83646 | 52.18 | 17.91 |
| 407 | 3.4141 | 3.2503 | 2468.432 | 42.83646 | 52.03 | 17.67 |
| 408 | 3.4143 | 3.254 | 2467.976 | 42.84102 | 52.4 | 18.24 |
| 409 | 3.4145 | 3.2548 | 2467.52 | 42.84557 | 52.48 | 18.36 |
| 410 | 3.4147 | 3.2559 | 2467.064 | 42.85013 | 52.59 | 18.52 |
| 411 | 3.4146 | 3.2572 | 2467.292 | 42.84785 | 52.72 | 18.73 |
| 412 | 3.4147 | 3.2569 | 2467.064 | 42.85013 | 52.69 | 18.68 |
| 413 | 3.4148 | 3.2578 | 2466.836 | 42.8524 | 52.78 | 18.81 |
| 414 | 3.4152 | 3.2577 | 2465.925 | 42.86151 | 52.77 | 18.78 |
| 415 | 3.4153 | 3.2577 | 2465.697 | 42.86379 | 52.77 | 18.77 |
| 416 | 3.4157 | 3.2584 | 2464.786 | 42.8729 | 52.84 | 18.86 |
| 417 | 3.4168 | 3.2582 | 2462.282 | 42.89797 | 52.82 | 18.78 |
| 418 | 3.4176 | 3.2603 | 2460.462 | 42.91621 | 53.03 | 19.07 |
| 419 | 3.4188 | 3.2604 | 2457.733 | 42.94359 | 53.04 | 19.04 |
| 420 | 3.4207 | 3.2607 | 2453.416 | 42.98696 | 53.07 | 19.00 |
| 421 | 3.4217 | 3.2599 | 2451.146 | 43.00981 | 52.99 | 18.83 |
| 422 | 3.4231 | 3.261 | 2447.97 | 43.04181 | 53.1 | 18.94 |
| 423 | 3.4266 | 3.2617 | 2440.042 | 43.12191 | 53.17 | 18.90 |
| 424 | 3.4281 | 3.2614 | 2436.649 | 43.15628 | 53.14 | 18.79 |
| 425 | 3.4301 | 3.2617 | 2432.13 | 43.20215 | 53.17 | 18.75 |
| 426 | 3.4322 | 3.262 | 2427.39 | 43.25036 | 53.2 | 18.70 |
| 427 | 3.4341 | 3.2633 | 2423.107 | 43.29402 | 53.33 | 18.82 |
| 428 | 3.4364 | 3.2635 | 2417.929 | 43.34692 | 53.35 | 18.75 |
| 429 | 3.4405 | 3.2636 | 2408.715 | 43.44138 | 53.36 | 18.59 |
| 430 | 3.4424 | 3.2641 | 2404.452 | 43.48522 | 53.41 | 18.58 |
| 431 | 3.445 | 3.2638 | 2398.627 | 43.54528 | 53.38 | 18.42 |
| 432 | 3.4477 | 3.2642 | 2392.587 | 43.60773 | 53.42 | 18.37 |
| 433 | 3.4501 | 3.2644 | 2387.226 | 43.66331 | 53.44 | 18.29 |
| 434 | 3.451 | 3.2643 | 2385.218 | 43.68417 | 53.43 | 18.24 |
| 435 | 3.4521 | 3.2645 | 2382.764 | 43.70968 | 53.45 | 18.22 |
| 436 | 3.4532 | 3.2644 | 2380.313 | 43.7352 | 53.44 | 18.16 |
| 437 | 3.455 | 3.2648 | 2376.304 | 43.777 | 53.48 | 18.14 |
| 438 | 3.4554 | 3.2648 | 2375.414 | 43.78629 | 53.48 | 18.13 |
| 439 | 3.4565 | 3.2649 | 2372.967 | 43.81186 | 53.49 | 18.09 |
| 440 | 3.4575 | 3.265 | 2370.743 | 43.83511 | 53.5 | 18.07 |
| 441 | 3.4578 | 3.265 | 2370.077 | 43.84209 | 53.5 | 18.05 |
| 442 | 3.4585 | 3.265 | 2368.521 | 43.85838 | 53.5 | 18.02 |
| 443 | 3.4589 | 3.265 | 2367.633 | 43.86769 | 53.5 | 18.00 |
| 444 | 3.4593 | 3.2651 | 2366.745 | 43.877 | 53.51 | 18.00 |
| 445 | 3.4596 | 3.2651 | 2366.079 | 43.88398 | 53.51 | 17.99 |
| 446 | 3.4604 | 3.2653 | 2364.303 | 43.90262 | 53.53 | 17.99 |
| 447 | 3.4607 | 3.2653 | 2363.637 | 43.9096 | 53.53 | 17.97 |
| 448 | 3.461 | 3.2652 | 2362.972 | 43.91659 | 53.52 | 17.94 |
| 449 | 3.4614 | 3.2653 | 2362.085 | 43.92592 | 53.53 | 17.94 |
| 450 | 3.4618 | 3.2652 | 2361.198 | 43.93524 | 53.52 | 17.91 |
| 451 | 3.4627 | 3.2653 | 2359.203 | 43.95622 | 53.53 | 17.88 |
| 452 | 3.4633 | 3.2653 | 2357.874 | 43.97022 | 53.53 | 17.86 |
| 453 | 3.4651 | 3.2653 | 2353.888 | 44.01223 | 53.53 | 17.78 |
| 454 | 3.465 | 3.2654 | 2354.11 | 44.00989 | 53.54 | 17.80 |
| 455 | 3.4668 | 3.2654 | 2350.128 | 44.05194 | 53.54 | 17.72 |
| 456 | 3.4674 | 3.2654 | 2348.802 | 44.06596 | 53.54 | 17.70 |
| 457 | 3.4689 | 3.2655 | 2345.489 | 44.10104 | 53.55 | 17.65 |
| 458 | 3.4704 | 3.2655 | 2342.178 | 44.13615 | 53.55 | 17.58 |
| 459 | 3.4705 | 3.2654 | 2341.957 | 44.13849 | 53.54 | 17.56 |
| 460 | 3.4728 | 3.2654 | 2336.887 | 44.19238 | 53.54 | 17.46 |
| 461 | 3.4742 | 3.2655 | 2333.804 | 44.22521 | 53.55 | 17.41 |
| 462 | 3.4766 | 3.2655 | 2328.524 | 44.28154 | 53.55 | 17.31 |
| 463 | 3.477 | 3.2657 | 2327.645 | 44.29094 | 53.57 | 17.32 |
| 464 | 3.4798 | 3.2657 | 2321.496 | 44.35677 | 53.57 | 17.20 |
| 465 | 3.4808 | 3.2656 | 2319.303 | 44.3803 | 53.56 | 17.14 |
| 466 | 3.4814 | 3.2658 | 2317.987 | 44.39443 | 53.58 | 17.14 |
| 467 | 3.4844 | 3.2658 | 2311.416 | 44.46512 | 53.58 | 17.01 |
| 468 | 3.4866 | 3.2657 | 2306.605 | 44.51703 | 53.57 | 16.90 |
| 469 | 3.489 | 3.2657 | 2301.363 | 44.57373 | 53.57 | 16.79 |
| 470 | 3.4897 | 3.2658 | 2299.835 | 44.59029 | 53.58 | 16.78 |
| 471 | 3.4909 | 3.266 | 2297.218 | 44.61867 | 53.6 | 16.76 |
| 472 | 3.4933 | 3.2659 | 2291.989 | 44.6755 | 53.59 | 16.63 |
| 473 | 3.4948 | 3.266 | 2288.724 | 44.71106 | 53.6 | 16.58 |
| 474 | 3.4956 | 3.2661 | 2286.984 | 44.73003 | 53.61 | 16.56 |
| 475 | 3.4961 | 3.2664 | 2285.897 | 44.74189 | 53.64 | 16.59 |
| 476 | 3.4986 | 3.2667 | 2280.466 | 44.80125 | 53.67 | 16.52 |
| 477 | 3.4977 | 3.2667 | 2282.421 | 44.77987 | 53.67 | 16.56 |
| 478 | 3.4991 | 3.2664 | 2279.381 | 44.81313 | 53.64 | 16.46 |
| 479 | 3.4998 | 3.2667 | 2277.862 | 44.82977 | 53.67 | 16.47 |
| 480 | 3.5008 | 3.2667 | 2275.694 | 44.85356 | 53.67 | 16.43 |
| 481 | 3.5014 | 3.2665 | 2274.393 | 44.86783 | 53.65 | 16.37 |
| 482 | 3.5021 | 3.2664 | 2272.876 | 44.88449 | 53.64 | 16.32 |
| 483 | 3.5025 | 3.267 | 2272.01 | 44.89401 | 53.7 | 16.40 |
| 484 | 3.5028 | 3.2669 | 2271.36 | 44.90116 | 53.69 | 16.37 |
| 485 | 3.5029 | 3.2669 | 2271.144 | 44.90354 | 53.69 | 16.37 |
| 486 | 3.5035 | 3.2672 | 2269.845 | 44.91783 | 53.72 | 16.39 |
| 487 | 3.5038 | 3.2677 | 2269.195 | 44.92498 | 53.77 | 16.45 |
| 488 | 3.504 | 3.2677 | 2268.763 | 44.92974 | 53.77 | 16.44 |
| 489 | 3.5041 | 3.2682 | 2268.546 | 44.93212 | 53.82 | 16.51 |
| 490 | 3.5045 | 3.2679 | 2267.681 | 44.94166 | 53.79 | 16.45 |
| 491 | 3.5046 | 3.2686 | 2267.464 | 44.94404 | 53.86 | 16.55 |
| 492 | 3.5046 | 3.269 | 2267.464 | 44.94404 | 53.9 | 16.62 |
| 493 | 3.5047 | 3.2692 | 2267.248 | 44.94642 | 53.92 | 16.64 |
| 494 | 3.5047 | 3.269 | 2267.248 | 44.94642 | 53.9 | 16.61 |
| 495 | 3.5049 | 3.2697 | 2266.815 | 44.95119 | 53.97 | 16.71 |
| 496 | 3.5054 | 3.2697 | 2265.734 | 44.96311 | 53.97 | 16.69 |
| 497 | 3.5054 | 3.2705 | 2265.734 | 44.96311 | 54.05 | 16.81 |
| 498 | 3.5059 | 3.2701 | 2264.653 | 44.97504 | 54.01 | 16.73 |
| 499 | 3.5069 | 3.2707 | 2262.492 | 44.9989 | 54.07 | 16.78 |
| 500 | 3.508 | 3.2708 | 2260.116 | 45.02516 | 54.08 | 16.74 |
| 501 | 3.5081 | 3.2715 | 2259.9 | 45.02754 | 54.15 | 16.85 |
| 502 | 3.5083 | 3.2716 | 2259.469 | 45.03232 | 54.16 | 16.85 |
| 503 | 3.5092 | 3.2721 | 2257.526 | 45.05382 | 54.21 | 16.89 |
| 504 | 3.5105 | 3.2726 | 2254.722 | 45.08489 | 54.26 | 16.91 |
| 505 | 3.5116 | 3.2722 | 2252.352 | 45.1112 | 54.22 | 16.80 |
| 506 | 3.5134 | 3.2737 | 2248.475 | 45.15429 | 54.37 | 16.95 |
| 507 | 3.5143 | 3.2731 | 2246.538 | 45.17585 | 54.31 | 16.82 |
| 508 | 3.5171 | 3.275 | 2240.519 | 45.24299 | 54.5 | 16.99 |
| 509 | 3.5167 | 3.2747 | 2241.379 | 45.23339 | 54.47 | 16.96 |
| 510 | 3.5181 | 3.2751 | 2238.372 | 45.26699 | 54.51 | 16.96 |
| 511 | 3.5226 | 3.2761 | 2228.724 | 45.37516 | 54.61 | 16.91 |
| 512 | 3.5243 | 3.2747 | 2225.086 | 45.4161 | 54.47 | 16.62 |
| 513 | 3.5261 | 3.2763 | 2221.237 | 45.45948 | 54.63 | 16.79 |
| 514 | 3.5271 | 3.2769 | 2219.101 | 45.4836 | 54.69 | 16.83 |
| 515 | 3.5278 | 3.2776 | 2217.606 | 45.5005 | 54.76 | 16.91 |
| 516 | 3.5313 | 3.2778 | 2210.141 | 45.58505 | 54.78 | 16.79 |
| 517 | 3.5337 | 3.2785 | 2205.031 | 45.64313 | 54.85 | 16.79 |
| 518 | 3.5352 | 3.2807 | 2201.841 | 45.67947 | 55.07 | 17.05 |
| 519 | 3.536 | 3.2803 | 2200.14 | 45.69886 | 55.03 | 16.96 |
| 520 | 3.5381 | 3.2801 | 2195.68 | 45.74981 | 55.01 | 16.83 |
| 521 | 3.5379 | 3.2803 | 2196.105 | 45.74495 | 55.03 | 16.87 |
| 522 | 3.5393 | 3.2804 | 2193.134 | 45.77894 | 55.04 | 16.83 |
| 523 | 3.5404 | 3.2805 | 2190.802 | 45.80567 | 55.05 | 16.79 |
| 524 | 3.5424 | 3.2807 | 2186.565 | 45.85431 | 55.07 | 16.73 |
| 525 | 3.5431 | 3.281 | 2185.083 | 45.87134 | 55.1 | 16.75 |
| 526 | 3.5443 | 3.282 | 2182.544 | 45.90056 | 55.2 | 16.85 |
| 527 | 3.5451 | 3.2842 | 2180.852 | 45.92005 | 55.42 | 17.14 |
| 528 | 3.5458 | 3.2853 | 2179.372 | 45.93711 | 55.53 | 17.28 |
| 529 | 3.5466 | 3.2867 | 2177.682 | 45.95662 | 55.67 | 17.45 |
| 530 | 3.5469 | 3.2869 | 2177.049 | 45.96394 | 55.69 | 17.46 |
| 531 | 3.5476 | 3.286 | 2175.57 | 45.98102 | 55.6 | 17.30 |
| 532 | 3.5476 | 3.2861 | 2175.57 | 45.98102 | 55.61 | 17.32 |
| 533 | 3.5481 | 3.2874 | 2174.515 | 45.99322 | 55.74 | 17.49 |
| 534 | 3.5481 | 3.286 | 2174.515 | 45.99322 | 55.6 | 17.28 |
| 535 | 3.5486 | 3.2876 | 2173.46 | 46.00542 | 55.76 | 17.49 |
| 536 | 3.5488 | 3.2884 | 2173.038 | 46.01031 | 55.84 | 17.60 |
| 537 | 3.5489 | 3.2886 | 2172.827 | 46.01275 | 55.86 | 17.63 |
| 538 | 3.549 | 3.2898 | 2172.616 | 46.01519 | 55.98 | 17.80 |
| 539 | 3.5491 | 3.29 | 2172.405 | 46.01763 | 56 | 17.83 |
| 540 | 3.5492 | 3.2916 | 2172.194 | 46.02008 | 56.16 | 18.06 |
| 541 | 3.5493 | 3.2925 | 2171.983 | 46.02252 | 56.25 | 18.18 |
| 542 | 3.5493 | 3.2918 | 2171.983 | 46.02252 | 56.18 | 18.08 |
| 543 | 3.5494 | 3.2927 | 2171.772 | 46.02496 | 56.27 | 18.21 |
| 544 | 3.5494 | 3.2926 | 2171.772 | 46.02496 | 56.26 | 18.19 |
| 545 | 3.5495 | 3.2936 | 2171.561 | 46.0274 | 56.36 | 18.33 |
| 546 | 3.5495 | 3.2931 | 2171.561 | 46.0274 | 56.31 | 18.26 |
| 547 | 3.5495 | 3.2945 | 2171.561 | 46.0274 | 56.45 | 18.46 |
| 548 | 3.5495 | 3.2955 | 2171.561 | 46.0274 | 56.55 | 18.61 |
| 549 | 3.5495 | 3.2975 | 2171.561 | 46.0274 | 56.75 | 18.89 |
| 550 | 3.5495 | 3.2974 | 2171.561 | 46.0274 | 56.74 | 18.88 |
| 551 | 3.5495 | 3.2972 | 2171.561 | 46.0274 | 56.72 | 18.85 |
| 552 | 3.5495 | 3.2985 | 2171.561 | 46.0274 | 56.85 | 19.04 |
| 553 | 3.5495 | 3.3003 | 2171.561 | 46.0274 | 57.03 | 19.29 |
| 554 | 3.5495 | 3.3002 | 2171.561 | 46.0274 | 57.02 | 19.28 |
| 555 | 3.5495 | 3.3005 | 2171.561 | 46.0274 | 57.05 | 19.32 |
| 556 | 3.5495 | 3.301 | 2171.561 | 46.0274 | 57.1 | 19.39 |
| 557 | 3.5495 | 3.301 | 2171.561 | 46.0274 | 57.1 | 19.39 |
| 558 | 3.5496 | 3.3014 | 2171.35 | 46.02985 | 57.14 | 19.44 |
| 559 | 3.5495 | 3.3008 | 2171.561 | 46.0274 | 57.08 | 19.36 |
| 560 | 3.5497 | 3.301 | 2171.14 | 46.03229 | 57.1 | 19.38 |
| 561 | 3.5496 | 3.3013 | 2171.35 | 46.02985 | 57.13 | 19.43 |
| 562 | 3.5497 | 3.3028 | 2171.14 | 46.03229 | 57.28 | 19.64 |
| 563 | 3.5498 | 3.3022 | 2170.929 | 46.03473 | 57.22 | 19.55 |
| 564 | 3.5498 | 3.3026 | 2170.929 | 46.03473 | 57.26 | 19.60 |
| 565 | 3.5499 | 3.3018 | 2170.718 | 46.03718 | 57.18 | 19.49 |
| 566 | 3.5499 | 3.3027 | 2170.718 | 46.03718 | 57.27 | 19.61 |
| 567 | 3.5501 | 3.3032 | 2170.296 | 46.04206 | 57.32 | 19.68 |
| 568 | 3.55 | 3.3035 | 2170.507 | 46.03962 | 57.35 | 19.72 |
| 569 | 3.5501 | 3.3038 | 2170.296 | 46.04206 | 57.38 | 19.76 |
| 570 | 3.5502 | 3.3036 | 2170.085 | 46.04451 | 57.36 | 19.73 |
| 571 | 3.5507 | 3.3029 | 2169.032 | 46.05673 | 57.29 | 19.61 |
| 572 | 3.5509 | 3.3028 | 2168.61 | 46.06162 | 57.28 | 19.59 |
| 573 | 3.5507 | 3.3035 | 2169.032 | 46.05673 | 57.35 | 19.69 |
| 574 | 3.5507 | 3.3042 | 2169.032 | 46.05673 | 57.42 | 19.79 |
| 575 | 3.5514 | 3.3056 | 2167.557 | 46.07384 | 57.56 | 19.96 |
| 576 | 3.5511 | 3.3057 | 2168.189 | 46.06651 | 57.57 | 19.98 |
| 577 | 3.552 | 3.3057 | 2166.293 | 46.08852 | 57.57 | 19.94 |
| 578 | 3.5521 | 3.3063 | 2166.082 | 46.09096 | 57.63 | 20.02 |
| 579 | 3.5522 | 3.3063 | 2165.872 | 46.09341 | 57.63 | 20.02 |
| 580 | 3.5528 | 3.3063 | 2164.608 | 46.10809 | 57.63 | 19.99 |
| 581 | 3.5535 | 3.3066 | 2163.135 | 46.12523 | 57.66 | 20.00 |
| 582 | 3.5531 | 3.3067 | 2163.977 | 46.11544 | 57.67 | 20.04 |
| 583 | 3.5546 | 3.3069 | 2160.821 | 46.15217 | 57.69 | 20.00 |
| 584 | 3.5554 | 3.3073 | 2159.139 | 46.17177 | 57.73 | 20.02 |
| 585 | 3.5555 | 3.307 | 2158.929 | 46.17422 | 57.7 | 19.98 |
| 586 | 3.5571 | 3.3076 | 2155.568 | 46.21346 | 57.76 | 19.99 |
| 587 | 3.5576 | 3.3075 | 2154.518 | 46.22573 | 57.75 | 19.96 |
| 588 | 3.558 | 3.3075 | 2153.678 | 46.23554 | 57.75 | 19.94 |
| 589 | 3.5592 | 3.3075 | 2151.161 | 46.26501 | 57.75 | 19.89 |
| 590 | 3.5593 | 3.3082 | 2150.951 | 46.26746 | 57.82 | 19.98 |
| 591 | 3.5608 | 3.3077 | 2147.806 | 46.30433 | 57.77 | 19.85 |
| 592 | 3.5601 | 3.3075 | 2149.274 | 46.28712 | 57.75 | 19.85 |
| 593 | 3.561 | 3.3077 | 2147.387 | 46.30924 | 57.77 | 19.84 |
| 594 | 3.5617 | 3.3082 | 2145.921 | 46.32646 | 57.82 | 19.88 |
| 595 | 3.563 | 3.3081 | 2143.199 | 46.35845 | 57.81 | 19.81 |
| 596 | 3.5652 | 3.3083 | 2138.597 | 46.41263 | 57.83 | 19.74 |
| 597 | 3.5662 | 3.3088 | 2136.508 | 46.43729 | 57.88 | 19.77 |
| 598 | 3.5691 | 3.3087 | 2130.454 | 46.50886 | 57.87 | 19.63 |
| 599 | 3.5713 | 3.3086 | 2125.868 | 46.56324 | 57.86 | 19.52 |