| LOCATION: Swansea | LLEOLIAD: abertawe |
| --- | --- |
| LEVEL 6 | LEFEL 6 |
| EXAMINATION PERIOD | CYFNOD ARHOLIADAU |
| may 2020 | mai 2020 |

MODULE CODE: ACDA5006

Transducers and Signal Conditioning

TIME:  2 HOURS

answer all questions

Equipment or materials required – Calculator

Adnoddau neu ddefnyddiau sydd eu hangen - cyfrifiannell

This is an open book exam

ANSWER ALL QUESTIONS

- The following table shows the output measurements of a voltmeter under two sets of conditions: used in an environment kept at 20°C, which is the temperature that it was calibrated at; and used in an environment at a temperature of 50°C

| Voltage readings at calibration temperature of 20°C (assumed correct) | Voltage readings at temperature of 50°C |
| --- | --- |
| 10.2 | 11.52 |
| 20.3 | 22.63 |
| 30.7 | 34.07 |
| 40.8 | 45.18 |

- Determine the zero drift when it is used in the 50°C environment, assuming that the measurement values when it was used in the 20°C environment are correct.

[3 Marks]

- Calculate the zero drift coefficient

[2 Marks]

- Determine the sensitivity drift when it is used in the 50°C environment, assuming that the measurement values when it was used in the 20°C environment are correct.

[3 Marks]

- Calculate the sensitivity drift coefficient

[2 Marks]

- The calibration of a pressure sensor with a range of 0-16 bar is checked by applying known inputs to it in steps of 0.2 bar.  The following measurements were obtained:

| Known input pressure (bar) | 0 | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Output reading (V) | 0.00 | 0.55 | 1.25 | 1.70 | 2.35 | 3.30 | 3.65 | 4.05 | 4.85 |

- Draw a graph by plotting the data points on the graph paper provided and draw a best-fit straight line through the data points. Use the graph to calculate:

[2 Marks]

- The measurement sensitivity of the pressure sensor

[4 Marks]

- The maximum non-linearity apparent in the measurements as a percentage of the full scale deflection.

[4 Marks]

- Explain the purpose of calibration. Include in your answer calibration frequency and environmental factors.

[10 Marks]

- A digital voltmeter has a resistance of 100kΩ and is connected in parallel with the 1MΩ resistor in figure Q4.

  - What would the true voltage across the 1MΩ resistor be if the voltmeter was not connected?

[2 Marks]

  - Assuming the voltmeter is calibrated correctly, what is the reading shown by the voltmeter?

[3 Marks]

  - State one method of making the voltmeter more accurate.

[5 Marks]

Figure Q4

- Describe the key specifications you would consider when selecting a flow sensor.  Compare the flow sensor requirements for:

  - Controlling the amount of chlorine added in a drinking water chlorination plant.

  - Domestic gas smart meter allowing the customer to see how much gas is being used in their home.

[10 Marks]

Diwedd/End