Shaun Farthing

Student number: 1708439

Abstract

Evaluation of current available weather stations, compare and contrast sensors, connectors and components required to produce a bespoke marine weather station. Produce full circuit schematic drawings and PCB design. Finishing with costings and evaluation.

Marine weather monitor/ logger

Module code: ACDA6002

Dr Gary Jones

Table of Contents

Tables and figures	2

Introduction	4

Literature review	5

Background research	5

Existing Systems	6

Sensor selection	8

Combined Temperature and Humidity	8

Wave height	10

Wind speed & direction	12

Sunlight brightness	14

GPS position	16

Connector selection	17

LED’s	19

Battery sizing & selection	20

Enclosure selection	21

Microcontroller selection	22

Non-volatile memory	24

Structure Diagram	25

Data acquisition cycle	26

Schematic design	27

Power circuit	27

Temperature & humidity	28

Wave height	28

Wind speed & direction	29

Sunlight brightness	29

GPS sensor	30

USB	30

Non-volatile memory	30

LED’s	31

PIC layout	32

PCB Design	33

Prototype build cost	37

Future Improvements	38

References	39

## Tables and figures

Figure 1 - Weatherpak marine	6

Figure 2 - RWS - 07	6

Figure 3 - Vantage Pro2	6

Figure 4 - HydroMet MAWS201	6

Figure 6 - Omega HX94BV1	9

Figure 7 - ATO - LEVS - ZP	11

Figure 8 – Gill Windsonic	13

Figure 9 - Apogee SP-215	15

Figure 10 - Standard horizon SCU-31	16

Figure 11 - 4 pin panel connector	18

Figure 12 - 4 pin cable connector	18

Figure 13 - 3 pin panel connector	18

Figure 14 - 3 pin cable connector	18

Figure 15 - 12v present LED	19

Figure 16 - Flashing indication	19

Figure 17 – Yuasa L35-90 battery	20

Figure 18 - Challenger TR5013 enclosure	21

Figure 19 - PIC18F25K50	23

Figure 20 - SanDisk SD Card	24

Figure 21 - Structure diagram	25

Figure 22 - Data acquisition cycle	26

Figure 23 - Proteus overall schematic design	27

Figure 24 - Power supply circuit	27

Figure 25 - Temperature & humidity	28

Figure 26 - Wave height	28

Figure 27 - Wind speed & direction	29

Figure 28 - Sunlight brightness	29

Figure 29 - GPS sensor circuit	30

Figure 30 - USB circuit	30

Figure 31 - Non-volatile memory circuit	30

Figure 32 - LED circuits	31

Figure 33 - PIC layout	32

Figure 34 - Easy-PC schematic	33

Figure 35 - Top copper layer of PCB	33

Figure 36 - Bottom copper layer	34

Figure 37 - PCB all layers	34

Figure 38 - Top layer 3D view	35

Figure 39 - Bottom layer 3D view	35

Figure 40 - Design rule check	36

Figure 41 - Prototype build cost	37

Table 1 - Existing systems	7

Table 2 - RH and temperature technology types	8

Table 3 - Manufacturer's RH & Temp	9

Table 4 - Wave height measurement technologies	10

Table 5 - Manufacturer's ultrasonic transducer	10

Table 6 - Wind speed & direction technology types	12

Table 7 – Manufacturer’s wind speed & direction sensors	13

Table 8 - Sunlight brightness technology types	14

Table 9 - Manufacturer's pyranometers	15

Table 10 - Manufacturers GPS antennas	16

Table 11 - 4 pin connectors	17

Table 12 - 3 pin connectors	18

Table 13 - Resolution requirements	22

Table 14 - PIC specification	23

Table 15 - Non-volatile memory requirement	24

## Introduction

This report details the tasks required to design a microcontroller based marine weather monitor/ logger. Firstly, current market systems are analysed, compared and evaluated. The report then looks at the different sensor types available to measure temperature, humidity, wave height, wind speed, wind direction, sunlight brightness and GPS position. Connector types and suitability for each sensor is then researched and documented. A microcontroller will be selected based on the chosen sensor requirements and calculations. Data storage, data transfer, power supply and a transportation system will be designed and documented in this report. The circuit design will be carried out using Proteus software, this will then transfer to the Easy-PC software for PCB design. The report concludes with a cost table for the prototype build and a full project review.

## Literature review

## Background research

Weather data acquisition is the process of sampling signals that measure real world physical conditions and converting the resulting samples into values that can be manipulated by a computer. [1] A marine weather monitor/ logger is a device that collects data related to the weather and environment using many different sensors. Weather station sensors may include air temperature, wind speed and direction, humidity and barometric pressure. [2]  Data acquisition systems serve a very important part of many systems, they tie together a range of sensors measuring lots of different processes. Many have the ability to visually relay live data and/ or store data to be transferred elsewhere later. [3]

## Existing Systems

Campbell scientific Weatherpak marine is a professional - level meteorological weather monitoring system in a compact, rugged, and easy-to-install package. It can be used in a wide range of industries and has the ability to be able to add additional instruments. [4]

Figure 1 - Weatherpak marine

Figure 2 - RWS - 07

Omni instruments RWS-07 is a portable, solid-state weather station. The RWS-07 can monitor pressure, temperature, rainfall and many other weather conditions. It has a 6-month battery capability and all data can be transferred in Excel or CSV format. [5]

Figure 3 - Vantage Pro2

The vantage pro 2 is a compact, robust and accurate weather station. It has multiple sensor options and offers output alarm contacts. Power can be provided by mains or battery. There is a hand-held visual console with the option of a weather link to a PC. [6]

The Vaisala HydroMet MAWS201 portable automatic weather station is designed for temporary installation. It’s lightweight, robust and suitable for demanding meteorological applications. The HydroMet is fully adjustable to suit uneven terrains. [7]

Figure 4 - HydroMet MAWS201

| Manufacturer | Campbell Scientific | Omni Instruments | Davis Instruments | Vaisala |
| --- | --- | --- | --- | --- |
| Model | Weatherpak Marine | RWS - 07 | Vantage Pro2 | HydroMet MAWS201 |
| Picture |  |  |  |  |
| Instrument types | Air temp, relative humidity, barometric pressure, wind speed/ direction, compass. | Wind Speed, direction, Temperature, Humidity, Rainfall Solar Radiation, Barometric Pressure | Wind speed, direction, rain collector, temperature, relative humidity, | Wind, pressure, air temperature, relative humidity, solar radiation, precipitation, soil/ water temperature, soil moisture |
| Input voltage | 9.6 – 16 vdc Optional power cable (AC/DC) | 12vdc | 5VDC adaptor with battery back up | 12VDC with solar panel and 7Ah battery back up |
| Inputs/ outputs | 6 integrated instruments plus additional through junction box | 4-8 x 4-20mA, 0-10v T/C, mV, 8 x Digital 2 x Counter | Integrated sensor suite (ISS) contains rain collector, anemometer, passive radiation shield. | 10 analogue inputs, 2 counter/ frequency inputs, internal channel for pressure sensor |
| Data transfer/ storage | NMEA 0183, RS-232, RS-485, TCP/IP, UHF, satellite | Direct RS232, USB | USB | RS-232, 2 wire RS-485, SDI-12 |
| Sample rate | 100Hz | 1 second to 12 hours | Dependant on sensor | Dependant on sensor |
| Logging duration/ battery life | Power supply | Up to 6 months | 1 month on battery | 12 months |
| Storage/ memory | 4MB | 1.8 million readings | Weatherlink option to upload straight to computer/ database. | 3.3MB internal, 2GB on Compact Flash card |
| Durability | Military-Standard 6061-T6 anti-corrosive aluminium; coated with a baked-on thermoplastic resin | IP67 weatherproof case | UV resistant ABS plastic | Anodized aluminium, NEMA 4X, IP66. |
| Cost | £2,740.70 | £2,414.80 | £909.90 | £2,987.00 |
| Web address | https://www.campbellsci.co.uk/weatherpak-marine [4] | https://www.omniinstruments.co.uk/rws-07-portable-weather-station-1082.html [5] | https://www.weather-station-products.co.uk/cabled-vantage-pro2-bundle [6] | https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/maws201 [7] |

Table 1 - Existing systems

## Sensor selection

# Combined Temperature and Humidity

| Sensor type | Humidity and Temperature integrated circuit sensor | Relative humidity (capacitive/ resistive) and temperature sensor ( Thermistor/ RTD/ TC temperature sensor) | Thermometer hygrometer |
| --- | --- | --- | --- |
| Cost | Low | Medium | High |
| Accuracy | High | High | High |
| Environmental suitability | Low | High | Low |
| Operating range suitability | High | Medium | Medium |
| Installation suitability | Low | Medium | High |

Table 2 - RH and temperature technology types

The sensor technology chosen is a RH and temperature transducer.

Manufacturer’s options;

| Manufacturer | Omega | Measurement Systems LTD | Omni instruments | Dwyer instruments |
| --- | --- | --- | --- | --- |
| Model | HX94BV1 | RHT-02 | Hygroclip 2. Mini probes | WHT-333 |
| Image |  |  |  |  |
| Power supply | 9–30 Vdc | 24V ac/dc | 3.2-5VDC | 10-35VDC |
| Power consumption | 100mA | 10mA | 4.5mA | 5mA |
| Output | 0–5vdc | 0-5Vdc | 0-1 V | 0-5V |
| Environmental protection | IP66 stainless steel NEMA4 | IP65 plastic | Nickel plated brass | IP54, NEMA3S, ABS. |
| Connections | 4 pin | 5 pin | 7 pins | 4 pins |
| Operating range | Temp = 0-100 degC. RH = 0-100%. | Temp= -20+50degC. RH=0-95%. | Temp= -40+60degC. RH=0-100%. | Temp=-40+60degC. RH=0-100%. |
| Cost | £243.00 | £75.00 | £477.00 | £300.00 |
| Web | https://www.omega.co.uk/pptst/HX94B.html | https://www.measurementsystems.co.uk/rht-3-o_outside_temperature__relative_humidity_transmitter | https://www.omniinstruments.co.uk/temperature-and-humidity/temperature-relative-humidity-probes/hygroclip-2-mini-probes.html#upsell-product-div | https://www.dwyer-inst.co.uk/Product/AirQuality/Humidity-TemperatureTransmitters/SeriesWHT#specs |

Table 3 - Manufacturer's RH & Temp

On evaluation, the HX94BV1 from Omega was chosen because it offers the best environmental protection. Due to the harsh marine conditions it will be operating in, it was the most desirable. [8]

Figure 6 - Omega HX94BV1

# Wave height

| Sensor type | Pressure based sensor | Ultrasonic | Capacitive level probes |
| --- | --- | --- | --- |
| Cost | Low | Medium | High |
| Accuracy | Medium | High | Medium |
| Environmental suitability | High | High | High |
| Operating range suitability | High | High | Low |
| Installation suitability | Medium | High | Medium |

Table 4 - Wave height measurement technologies

The technology type chosen to measure wave height is ultrasonic.

Manufacturer’s options;

| Manufacturer | Siemens | ATO | APG | Gems sensors |
| --- | --- | --- | --- | --- |
| Model | LU240 | ATO-LEVS-ZP | IRU-9429-C6 | UCL-520 |
| Image |  |  |  |  |
| Power supply | 10.5-30 V DC | 12-24VDC | 12-28VDC | 12-28 VDC |
| Power consumption | 20mA | 125mA | 75mA | 22mA |
| Output | 4-20mA | 0-5V | 0-5V | 4-20mA |
| Environmental protection | Type 4X, Type 6, IP66 | IP68 | NEMA12 | NEMA4X, IP65 |
| Connections | 2 pin | 3 pin | 5 pin | 2 pin |
| Operating range | 0.2-12m | 0-40m | 0.15-10.67m | 1.8-8m |
| Cost | £984.20 | £536.50 | £499.40 | £1,006.87 |
| Web | https://www.finecontrols.co.uk/level-control/level-transmitters/ultrasonic-level-transmitters/ultrasonic-level-transmitter-lu240 | https://www.ato.com/ultrasonic-level-sensor-40m | https://www.bellflowsystems.co.uk/iru-9400-general-purpose-ultrasonic-level-transmitter.html?category_id=1210 | https://www.gemssensors.co.uk/level/continuous-transmitters/ultrasonic/ucl-520 |

Table 5 - Manufacturer's ultrasonic transducer

Following evaluation, the ATO – LEVS – ZP has been chosen, because it has a suitable IP rating, a favourable voltage output and is mid-range priced. [9]

Figure 7 - ATO - LEVS - ZP

Please note, when using the ultrasonic transducers to measure wave height, two transducers need to be used. One needs to be connected to an open ended pipe, this will allow the pipe to fill quickly, the second needs to be connected to a pipe with a small opening, and this will allow a damped signal and slow changing level. The difference between the two readings is the wave height.

# Wind speed & direction

| Sensor type | Ultrasonic | Vane anemometer | Cup anemometer | Thermal anemometer |
| --- | --- | --- | --- | --- |
| Cost | High | Medium | Low | High |
| Accuracy | High | High | High | High |
| Environmental suitability | High | Medium | Low | Medium |
| Operating range suitability | High | High | Low | High |
| Installation suitability | High | High | Medium | Medium |

Table 6 - Wind speed & direction technology types

The technology type chosen is ultrasonic as there are no moving parts. All other options have physical, mechanical moving parts. The less moving parts, the less susceptible to damage the instrument will be.

Manufacturer’s options;

| Manufacturer | Nova Lynx | Richard Paul Russell | FT Technologies | Gill |
| --- | --- | --- | --- | --- |
| Model | 200-8100 | WS3 | FT742 | Windsonic |
| Image |  |  |  |  |
| Power supply | 12-30 Vdc | 7-30 V dc | 6-30V DC | 5-30 V DC |
| Power consumption | 333mA | 5.5mA | 29mA | 44mA |
| Output | 0-5V | 0-5V | RS422, RS485 | 0-5V |
| Environmental protection | UV stabilized thermoplastic, Stainless steel and anodised aluminium. | IP65 | IP66 | IP66 |
| Connections | 5 pin | 5 pin | 8 pin | 5 pin |
| Operating range | Wind speed = 0-40m/s. Wind direction = 0-360 | Wind speed = 0-60m/s. Wind direction = 0-359 | Wind speed = 0-75m/s. Wind direction = 0-360 | Wind speed = 0-60m/s. Wind direction = 0-359 |
| Cost | £1950.00 | £1272.00 | £1250.00 | £640.00 |
| Web | https://novalynx.com/store/pc/200-81000-Ultrasonic-Anemometer-19p335.htm | https://www.r-p-r.co.uk/windsystems/windsonic.php | https://fttechnologies.com/wind-sensors/ft7-series/surface-mount/ | https://www.omniinstruments.co.uk/weather-stations-and-instruments/wind-speed-and-direction-sensors/windsonic-solid-state-wind-speed-and-direct [10]ion-sensor.html#upsell-product-div |

Table 7 – Manufacturer’s wind speed & direction sensors

Following comparison, the chosen instrument is the Windsonic by Gill as it’s most cost effective. All instruments were suitable. [10]

Figure 8 – Gill Windsonic

# Sunlight brightness

| Sensor type | Pyranometer | Pyrheliometers | Light dependant resistor | Spectrometer |
| --- | --- | --- | --- | --- |
| Cost | Medium | High | Low | High |
| Accuracy | High | High | Medium | Medium |
| Environmental suitability | High | High | Medium | Low |
| Operating range suitability | High | High | Medium | High |
| Installation suitability | High | High | Low | Low |

Table 8 - Sunlight brightness technology types

The sensor type selected is a pyranometer. The pyranometer offered the most suitable all round option with a mid-range price.

Manufacturer’s options;

| Manufacturer | Kipp & Zonen | Hukesflux | Apogee Instruments | Li-cor |
| --- | --- | --- | --- | --- |
| Model | SP Lite2 | SR05-D1A3 | SP-215 | LI-200R |
| Image |  |  |  |  |
| Power supply |  | 5-30VDC | 5.5-24 V DC | 3.8-28 V |
| Power consumption |  | 6mA | 0.3mA | 1mA |
| Output |  | 0-1V | 0-5V | -2.5-5V |
| Environmental protection | IP67 | IP67 | IP68 marine grade stainless steel | Weatherproof anodised aluminium |
| Connections | 3 pin | 4 pin | 4 pin | 3 pin |
| Operating range | 400-100nm | 285-3000nm | 360-1120nm | 400-1100nm |
| Cost | £330.00 |  | £197.00 | £325.00 |
| Web | https://www.omniinstruments.co.uk/weather-stations-and-instruments/pyranometers-solar-irradiance/sp-lite2-silicon-pyranometer.html#upsell-product-div | https://www.hukseflux.com/products/solar-radiation-sensors/pyranometers/sr05-d1a3-pyranometer | https://www.apogeeinstruments.com/sp-215-ss-amplified-0-5-volt-pyranometer/#product-tab-description | https://www.licor.com/env/products/light/pyranometer.html?gclid=EAIaIQobChMI8cjgpqmP5wIVFODtCh3klgvDEAMYAiAAEgJ3ePD_BwE |

Table 9 - Manufacturer's pyranometers

After comparing, the Apogee SP-215 is chosen. It offers the best environmental protection, a desired output and is most reasonably priced. [11]

Figure 9 - Apogee SP-215

# GPS position

| Manufacturer | Garmin | B&G | SparkFun | Standard Horizon |
| --- | --- | --- | --- | --- |
| Model | GPS16X-HVS | ZG100 | GPS-12751 | SCU-31 |
| Image |  |  |  |  |
| Power supply | 8-40 Vdc | 9-18v | 4.5-6.5Vdc | 10-35 VDC |
| Power consumption | 65mA | 100mA | 55mA | 25mA |
| Output | Position <15m | Position 3m | Position <2.5m | Position 2.5m |
| Environmental protection | IPX7, immersed in 1m of water for 30m | IPX-7 | Requires suitable enclosure | IPX7 |
| Connections | 8 pins |  | 6 | 4 pin |
| Cost | £87.00 | £268.00 | £32.00 | £109.00 |
| Web | https://www.campbellsci.com/gps16x-hvs | https://www.bandg.com/bg/type/instrument-sensors-and-transducers/gps-sensors/gps-antenna-bg-zg100-module-pack/#prl_product_details | https://www.mouser.co.uk/ProductDetail/SparkFun/GPS-12751?qs=sGAEpiMZZMsjLMBIknjmkr7UcSbxhsyWkbeQM%2FV612RTadYIA5zSIw%3D%3D | https://www.cactusnav.com/standard-horizon-external-antenna-with-cable-p-27129.html?gclid=EAIaIQobChMIg-m95sWP5wIVArTtCh2OMQ-tEAQYASABEgIRUPD_BwE |

Table 10 - Manufacturers GPS antennas

After comparing, the Standard horizon SCU-31 was chosen. It boasts a high accuracy position, suitable IP rating and minimal current draw. [12]

Figure 10 - Standard horizon SCU-31

# Connector selection

All connectors are to have a high IP rating (IP67/68) to protect against water ingress, gold plated contacts to prevent corrosion and possible bad connections and a weather resistant body material.

Temperature & humidity – 4 contacts

Pin 1 - + power supply, pin 2 – 0v, pin 3 – temperature output, and pin 4 – humidity output.

Wave height – 3 contacts

Pin 1 - + power supply, pin 2 – 0v, pin 3 – signal.

Wind speed & direction – 4 contacts

Pin 1 – signal ground, pin 2 - + power supply, pin 3 – wind direction, pin 4 – wind speed.

Sunlight brightness – 4 contacts

Pin 1 - + power supply, pin 2 – ground, pin 3 – signal, and pin 4 – shield/ ground.

GPS position – 4 contacts

Pin 1 - + power supply, pin 2 – ground, pin 3 – R/X signal, and pin 4 – T/X signal.

All connectors require 3 or 4 pins. Research has been completed to determine the most suitable connector, the selected connector has been highlighted in yellow on the tables below. Both the connector types selected meet the IP rating required, have gold plated, non-corrosive contacts and also a weather resistant body. The same connectors are going to be used on all sensors requiring 4 pins. The sensor and associated connector will both be identified using a coloured marker to avoid plugging the wrong connector into the wrong sensor.

| Manufacturer | Amphenol | Amphenol | Souriau-sunbank | Souriau-sunbank | Lumberg | Lumberg |
| --- | --- | --- | --- | --- | --- | --- |
| Model No. | C09131H0041002 | C091-31N004-1002 | UTS7104S | UTS6JC124P | 030404 | 033204 |
| Supplier | RS Components – 157 - 2461 | RS Components – 157-2534 | Digi-key SOU1280-ND | Digi-key SOU1300-ND | RS Components – 533-7795 | RS Components – 170-5486 |
| Image |  |  |  |  |  |  |
| Contacts | 4 | 4 | 4 | 4 | 4 | 4 |
| IP rating | IP67 | IP67 | IP69K | IP69K | IP68 | IP68 |
| Contact material | Silver plated | Silver plated | Gold plated | Gold plated | Gold plated | Gold plated |
| Body material | Die cast nickel plated | Die cast aluminium, nickel plated | Thermoplastic | Thermoplastic | Glass filled polyamide (PA GF) | Glass filled polyamide (PA GF) |
| Cost | £9.82 | £6.41 | £8.88 | £13.55 | £4.52 | £5.34 |

Table 11 - 4 pin connectors

| Manufacturer | TE Connectivity | TE Connectivity | MBS Connect | MBS Connect | Bulgin | Bulgin |
| --- | --- | --- | --- | --- | --- | --- |
| Model No. | HD10-3-96P | HD14-3-96P | SP1312/S3C | SP1310/S3IIC | PX0731/P | PX0730/S |
| Supplier | Digi-key 1734-1308-ND | Digi-key 1734-1300-ND | Farnell – CN21343 | Farnell – CN21315 | Farnell – CN08568 | Farnell – CN08589 |
| Image |  |  |  |  |  |  |
| Contacts | 3 | 3 | 3 | 3 | 3 | 3 |
| IP rating | IP68 | IP68 | IP68 | IP68 | IP68 | IP68 |
| Contact material | Gold plated | Gold plated | Gold plated | Gold plated | Nickel | Nickel |
| Body material | Glass filled polyamide (PA GF), nylon | Glass filled polyamide (PA GF), nylon | Nylon polyamide | Nylon polyamide | Nylon polyamide | Nylon polyamide |
| Cost | £4.38 | £4.74 | £2.43 | £2.43 | £8.12 | £6.39 |

Table 12 - 3 pin connectors

[13] [14] [15]

Figure 11 - 4 pin panel connector

Figure 12 - 4 pin cable connector

Figure 13 - 3 pin panel connector

Figure 14 - 3 pin cable connector

# LED’s

LED 1 indicates that 12v is present. LED 2 flashes once every 10 seconds, the timing for the flashing LED will be built into the software.

12v supply red LED: 5mm, forward voltage = 1.8v, forward current = 20mA.

Figure 15 - 12v present LED

Flashing once every 10 seconds green LED: 5mm, forward voltage = 2.1v, forward current = 20mA.

Figure 16 - Flashing indication

# Battery sizing & selection

| Component | Quantity | Power consumption | Total power consumption |
| --- | --- | --- | --- |
| Temperature & Humidity sensor | 1 | 100mA | 100mA |
| Wave height ultrasonic sensor | 2 | 125mA | 250mA |
| Wind speed & direction sensor | 1 | 44mA | 44mA |
| Sunlight brightness pyranometer | 1 | 0.3mA | 0.3mA |
| GPS position | 1 | 25mA | 25mA |
| LED’s | 2 | 20mA | 40mA |
| Total |  |  | 459.3mA |

Total power consumption is 459.3mA. The specification states that data should be acquired for a 7 day cycle (168 hours).

Total power consumption is multiplied by 168 to give the total power consumption over the monitoring period, 77,162.4mA (77.16A).

The battery chosen is a 12v, 90Ah, sealed, lead acid battery. Dimensions = 350mm x 174mm x 224mm. Weight = 23.6kg. [16]

Figure 17 – Yuasa L35-90 battery

# Enclosure selection

The enclosure selected is the Challenger TR5013, waterproof, wheeled travel case. IP67 rated and made from high strength moulded plastic. Measures 626mm x 492mm x 318mm which is sufficient to house all components and battery. The case has foam inserts which can be altered to the required shape to ensure there are no loose components. [17]

Figure 18 - Challenger TR5013 enclosure

# Microcontroller selection

To select a suitable microcontroller, we must consider, analogue and digital pin requirement, resolution, program memory requirement, RAM requirement, clock speed, and other connections such as USB and SD card.

Resolution calculations;

Temperature, logged every 0.5 degC change. Instrument range, 0-100 degC. 100÷0.5 = 200. 8 bits required.

Humidity, logged every 0.5% change. Instrument range, 0-100 degC. 100÷0.5 = 200. 8 bits required.

Wave height, logged every 50mm change. Instrument range, 0-40m. 40000÷50=800. 10 bits required.

Wind speed, logged every 0.5mph change. Instrument range, 0-135mph. 135÷0.5=270. 10 bits required.

Wind direction, logged every 0.5°. Instrument range, 0-359°. 359÷0.5=718. 10 bits required.

Sunlight brightness, logged every 5W/m2. Instrument range, 0-2000w/m. 2000÷5=400. 10 bits required.

| Signal | Pins | Analogue/ Digital | Log frequency | Resolution |
| --- | --- | --- | --- | --- |
| Temperature | 1 | Analogue | 1 s | 8 bit |
| Humidity | 1 | Analogue | 1 s | 8 bit |
| Wave height open | 1 | Analogue | 0.5 s | 10 bit |
| Wave height damp | 1 | Analogue | 0.5 s | 10 bit |
| Wind speed | 1 | Analogue | 0.5 s | 10 bit |
| Wind direction | 1 | Analogue | 0.5 s | 10 bit |
| Sunlight brightness | 1 | Analogue | 1 s | 10 bit |
| GPS position | 2 | Digital | 1 s | 8 bit |
| SD card | 4 | Digital |  |  |
| Start button | 1 | Digital |  |  |
| Tell | 1 | Digital |  |  |
| USB | 2 | D+/D- |  |  |
| Power | 3 |  |  |  |
| Total | 20 |  |  |  |

Table 13 - Resolution requirements

Microcontroller criteria.

| Voltage | 5v device |
| --- | --- |
| Digital signals | 8 bits |
| Analogue signals | 8 @ 10 bit resolution |
| USB | 2 bits |
| UART | Yes (1) |
| SPI | Yes (1) |
| Program memory | 16KB |
| RAM | 0.5KB |
| Clock speed | 16MHz |

Table 14 - PIC specification

Requirements were entered into microchip.com ‘parameter chart search’ function. The search narrowed down to a choice of 2 microcontrollers. PIC18F25K50 and PIC18F25K40. PIC18F25K50 was chosen over PIC18F25K40 as it has everything that is required and was most cost effective. [18]

Figure 19 - PIC18F25K50

# Non-volatile memory

| Signal | Signal type | Log frequency | Bytes | Total bytes |
| --- | --- | --- | --- | --- |
| Temperature | Analogue | 1 s | 1 | 1 |
| Humidity | Analogue | 1 s | 1 | 1 |
| Wave height open | Analogue | 0.5 s | 2 | 4 |
| Wave height damp | Analogue | 0.5 s | 2 | 4 |
| Wind speed | Analogue | 0.5 s | 2 | 4 |
| Wind direction | Analogue | 0.5 s | 2 | 4 |
| Sunlight brightness | Analogue | 1 s | 2 | 2 |
| GPS position | Digital | 1 s | 1 | 1 |

Table 15 - Non-volatile memory requirement

Total number of bytes is 21/ s.

21 x 60 secs = 1260 bytes/ min

1260 x 60 mins = 75.6KB/ hour

75.6 x 24 hours = 1814.4 KB (1.81MB)/ day

1.81 x 7 days = 12.67MB / 7 days

Total storage required for the 7-day monitoring is 12. 67MB.

The SD card chosen is 1GB by SanDisk, this will provide ample storage space for the application. [19]

Figure 20 - SanDisk SD Card

# Structure Diagram

Figure 21 - Structure diagram

# Data acquisition cycle

Figure 22 - Data acquisition cycle

# Schematic design

Proteus schematic.

Figure 23 - Proteus overall schematic design

# Power circuit

There is a 12v battery in the circuit. Immediately after the battery is a switch, when operated, puts 12v onto all sensors requiring a 12v supply, voltage regulator and to an LED to indicate that there is 12v present. The voltage regulator has a 5v output. From the voltage regulator, all components requiring a 5v supply and fed. Capacitor C9 acts as a filter and capacitor C10 acts as a tank. R8 is a de-coupler. There is also a button that sends a signal to the PIC and brings on a green LED that flashes every 10 seconds.

Figure 24 - Power supply circuit

Sensors used for temperature, humidity, wave height, wind speed/ direction and sunlight brightness all require a 12v supply. They also all provide a 0-5v output. All these circuits have a clamp, the clamp consists of a capacitor and resistor that form a filter and a Zener diode to protect against overvoltage.

# Temperature & humidity

Figure 25 - Temperature & humidity

# Wave height

Figure 26 - Wave height

# Wind speed & direction

Figure 27 - Wind speed & direction

# Sunlight brightness

Figure 28 - Sunlight brightness

# GPS sensor

Figure 29 - GPS sensor circuit

# USB

The USB is required to enable the device to connect to a computer. The VCC will be supplied from the computer. The D+/- are bi-directional pins for data transfer.

Figure 30 - USB circuit

# Non-volatile memory

The SD card has a digital input from the PIC, a digital output to the PIC, a clock rate signal to allow the PIC to synchronise and a chip select signal so that the PIC can identify the device to write to as it’s a BUS system, although, it is the only device on the BUS, the CS is connected.

Figure 31 - Non-volatile memory circuit

# LED’s

Resistor sizing calculations;

12v on red LED.

Vres=12v-1.8v=10.2 v

R=10.2 v0.020 A=510Ω

5v flashing green LED.

Vres=5v-2.1v=2.9 v

R=2.9  v0.020 A=145Ω

Figure 32 - LED circuits

# PIC layout

Figure 33 - PIC layout

# PCB Design

Easy-PC schematic.

Figure 34 - Easy-PC schematic

Top layer translated to PCB. Using the optimise nets function, the PCB tracks were drawn onto the board.

Figure 35 - Top copper layer of PCB

Dark blue, 12v copper. Light blue, 0v copper.

Figure 36 - Bottom copper layer

Figure 37 - PCB all layers

3D views.

Figure 38 - Top layer 3D view

Figure 39 - Bottom layer 3D view

Design rule check.

Figure 40 - Design rule check

# Prototype build cost

Figure 41 - Prototype build cost

# Future Improvements

- Additional weather monitoring e.g. air pressure, rain gauges, dew point.

- Bluetooth & Wi-Fi connectivity.

- Decrease system physical size e.g. use PCB mounted GPS sensor.

- Select sensors with smaller energy usage to minimise size and weight of battery requirement.

- Reduce cost by using more cost-effective sensors and connectors.

# References

| [1] | A. P. P. M. Suvarna K. Gosav, “Weather Data Acquisition System,” June 2015. [Online]. Available: https://www.ijraset.com/fileserve.php?FID=2702. [Accessed 13 January 2020]. |
| --- | --- |
| [2] | “What is a weather station?,” 13 January 2020. [Online]. Available: https://www.acurite.com/. |
| [3] | “Data Aquisition,” Omega, [Online]. Available: https://www.omega.co.uk/prodinfo/dataacquisition.html. [Accessed 13 January 2020]. |
| [4] | “Campbell Scientific,” [Online]. Available: https://www.campbellsci.com/weatherpak-marine. [Accessed 25 01 2020]. |
| [5] | “Omni Instruments,” [Online]. Available: https://www.omniinstruments.co.uk/rws-07-portable-weather-station-1082.html. [Accessed 15 01 2020]. |
| [6] | “Weather station products,” [Online]. Available: https://www.weather-station-products.co.uk/cabled-vantage-pro2-bundle. [Accessed 10 01 2020]. |
| [7] | “Vaisala,” [Online]. Available: https://www.vaisala.com/en/products/instruments-sensors-and-other-measurement-devices/weather-stations-and-sensors/maws201. [Accessed 07 01 2020]. |
| [8] | “Omega,” [Online]. Available: https://www.omega.co.uk/pptst/HX94B.html. [Accessed 12 01 2020]. |
| [9] | “ATO,” [Online]. Available: https://www.ato.com/ultrasonic-level-sensor-40m. [Accessed 05 01 2020]. |
| [10] | “Omni instruments,” [Online]. Available: https://www.omniinstruments.co.uk/weather-stations-and-instruments/wind-speed-and-direction-sensors/windsonic-solid-state-wind-speed-and-direction-sensor.html#upsell-product-div. [Accessed 23 12 2019]. |
| [11] | “Apogee,” [Online]. Available: https://www.apogeeinstruments.com/sp-215-ss-amplified-0-5-volt-pyranometer/#product-tab-description. [Accessed 22 01 2020]. |
| [12] | “Cactus nav,” [Online]. Available: https://www.cactusnav.com/standard-horizon-external-antenna-with-cable-p-27129.html?gclid=EAIaIQobChMIg-m95sWP5wIVArTtCh2OMQ-tEAQYASABEgIRUPD_BwE. [Accessed 22 12 2019]. |
| [13] | “RS Components,” [Online]. Available: https://uk.rs-online.com/web/. [Accessed 24 01 2020]. |
| [14] | “Digi Key,” [Online]. Available: https://www.digikey.com/. [Accessed 25 01 2020]. |
| [15] | “Farnell,” [Online]. Available: https://uk.farnell.com/. [Accessed 25 01 2020]. |
| [16] | [Online]. Available: https://www.tayna.co.uk/leisure-batteries/yuasa/l35-90/. [Accessed 21 01 2020]. |
| [17] | [Online]. Available: https://www.trifibre.co.uk/product/waterproof-tr5013-equipment-hard-carry-photography-drone-wheeled-travel-flight-case/. [Accessed 21 01 2020]. |
| [18] | “Microchip,” [Online]. Available: https://www.microchip.com/. [Accessed 25 01 2020]. |
| [19] | “Amazon,” [Online]. Available: https://www.amazon.co.uk/dp/B000EYAEK8/ref=psdc_296791011_t3_B00065KA7C. [Accessed 25 01 2020]. |