| LOCATION: Swansea | LLEOLIAD: abertawe |
| --- | --- |
| LEVEL 5 | LEFEL 5 |
| EXAMINATION PERIOD | CYFNOD ARHOLIADAU |
| JUly 2025 | Gorffennaf 2025 |

MODULE CODE: ACDA5006

Transducers and Signal Conditioning

TIME:  2 HOURS

answer all questions

Equipment or materials required – Calculator, Graph Paper

Adnoddau neu ddefnyddiau sydd eu hangen - Dim

ANSWER ALL QUESTIONS

- Define the following terms:

  - Transducer

[2 Marks]

  - Accuracy

[2 Marks]

  - Precision

[2 Marks]

  - Repeatability

[2 Marks]

  - Reproducibility

[2 Marks]

  - Range

[2 Marks]

  - Nonlinearity

[2 Marks]

  - Resolution

[2 Marks]

  - Sensitivity

[2 Marks]

  - Hysteresis

[2 Marks]

- For the circuit shown in Fig Q2:

  - Describe the purpose of the circuit.

[2 Marks]

  - Explain the operation of circuit and the function of each component

[8 Marks]

Fig Q2

- The calibration of a pressure sensor with a range of 0-16 bar is checked by applying known inputs to it in steps of 2 bar.  The following measurements were obtained:

| Known input pressure (bar) | 0 | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Output reading (V) | 0.00 | 0.55 | 1.25 | 1.70 | 2.35 | 3.30 | 3.65 | 4.05 | 4.85 |

- Draw a graph by plotting the data points on the graph paper provided and draw a best-fit straight line through the data points. Use the graph to calculate:

[2 Marks]

- The measurement sensitivity of the pressure sensor

[4 Marks]

- The maximum non-linearity apparent in the measurements as a percentage of the full scale deflection.

[4 Marks]

- Describe the key specifications you would consider when selecting a temperature sensor.  Compare the sensor requirements for:

  - Controlling the temperature of a commercial bakery oven.

  - Monitoring the coolant temperature in a petrol engine.

[10 Marks]

Diwedd/End