OME Major Project Proposal: Development of a Method for Propellant Stabiliser Quantitation using High Performance Liquid Chromatography – Triple Quadrupole Mass Spectrometry

BEng(Hons) Ordnance, Munitions and Explosives

ACCA6020

Quinn Dunbar

2217666

June 2026

Table of Contents

1.	Introduction	3

1.1.	Aims and Deliverables	3

2.	Literature Review	3

3.	Project Design	4

3.1.	Method Development	4

3.2.	Method Validation	4

3.3.	Data collection and analysis	4

4.	Resourcing	4

4.1.	Staff Resourcing	4

4.2.	Other costs	5

4.3.	Time planning	5

References	6

ANNEX A – UWTSD Ethics Form	8

Table of Tables

Table 1 - Time Commitments and Costings for staffing on this project	5

Table of Figures

Figure 1 - Gantt Chart of Projected timetable of project	5

# Introduction

Propellant formulations contain a number of stabiliser compounds that slow the degradation of the nitrate esters used in single, double and triple base propellants. The stabilisers deplete over time as they react with degradation products. Therefore, quantification of the amount of stabiliser present in a propellant sample is important to determine if the sample is safe to store and coupled with artificial aging can be used to estimate the shelf storage life of a sample [1].  This project will concern the development of a method using Ultra High Pressure Liquid Chromatography coupled triple quadrupole tandem Mass Spectrometry (UHPLC-MS/MS) for the separation, identification and quantification of stabilisers within a propellant sample. The method will then be validated to show that it is fit for purpose and operates within an acceptable working range and uncertainty [2].

## Aims and Deliverables

The aims for this project to deliver by the end of Financial Year 26/27 are:

- Develop a written method for the quantification of stabilisers within a propellant sample via UHPLC-MS/MS

- Produce a validation report covering the method showing it to be suitably robust and with acceptable working range.

# Literature Review

Nitrate esters such as nitrocellulose (NC) and nitroglycerin (NG)  naturally degrade over time to produce various decomposition products such as nitrogen oxide gases, some of which can catalyse further decomposition leading to a runaway autocatalytic reaction that can result in fire or unintended initiation. To allow NC-based propellants to be safe to store and use over a period of years stabiliser compounds are added to the formulation that react with the degradation products, slowing the rate of decomposition substantially [3]. To determine if a propellant is safe to store and if it will remain so over time the quantity of stabiliser present can be measured. This can be used in conjunction with artificial aging and the quantities before and after compared to determine if the stabiliser quantity remains within safe levels [4].

Many methods for the determination of the stability of propellants are presented in literature, including qualitative, quantitative and analytical methods [5]. UHPLC-MS/MS allows for high selectivity and sensitivity compared to other methods, such as HPLC-UV, through the use of a triple quadrupole mass spectrometer (QqQ MS). This is because a QqQ MS can perform Multiple Reaction Monitoring (MRM), a selective detection technique that captures ions of a desired mass, fragments the ions and then detects the daughter ions, with only selected transitions being measured, allowing for the elimination of any interferent effects, such as multiple components co-eluting [6].

The major reported use of HPLC-MS/MS analysis of stabilisers in literature is in the characterisation of Organic Gunshot Residues (OGSR) [7] [8] [9] [10] [11] [12]. These are a mixture of organic compounds formed when a firearm is discharged and deposited on nearby surfaces, such as the hands of the firer. A key identifying component of OGSR is the presence of stabiliser compounds that do not have any other common sources [13].

Laza et al. have published a validated method for the quantification of stabilisers in OGSR [14]. Using HPLC-MS/MS with electrospray ionisation they quantify the stabilisers Diphenylamine and its nitrated derivatives, Akardite II, methyl- and ethyl-centralites, and achieve limits of detection as low as 0.07 nmol L-1. This paper shall form the basis for the method development performed in this project.

# Project Design

## Method Development

Method development will cover three main phases: determination of UHPLC conditions, determination of MS conditions and testing of samples. UHPLC conditions include the eluent, flow rate and column chemistry, and will be identified using standards of the target stabilisers. MS conditions include determining the parent and daughter ions that will be used to identify each analyte, and the MS conditions that best promote these pairings, this will also be conducted using standard solutions. The final step of method development will be running propellant samples using the determined conditions to ensure that the method is compatible with real-world matrices.

## Method Validation

Method validation will be carried out to show the method is robust with regards to possible real-world variation. It will also allow the identification of Limits of Detection and Quantification (LoD and LoQ), along with demonstrating the linear working range and allowing for the calculation of the expanded uncertainty in the results.  Validation experiments will be designed that cover changes in the column, analyst and concentration of the samples, as well as any effects that the different sample matrices may have on the accuracy of the results.

## Data collection and analysis

The UHPLC-MS/MS instrument is capable of outputting data about the analytes under test in a variety of ways including graphical and numerical data, such as chromatograms, mass spectra, UV spectra, total ion chromatograms, and calculated concentration. These outputs will be reviewed during the method development and the most appropriate will be incorporated into the report generated by the method. During validation these figures will undergo statistical analysis to determine variance within the method and ensure it is accounted for through expanded uncertainty calculations.

# Resourcing

## Staff Resourcing

The primary staffing for this project is one apprentice, who will carry out the method development experiments. For the validation experiments they will be supported by six additional validation scientists (VS) to cover analyst to analyst variation. Output from the project will be reported to and supported by a Project Technical Authority (PTA), Lead Technical Reviewer (LTR) and Quality Manager (QM) who will provide technical guidance and review output for suitability. Time allocation and associated costs are detailed in Table 1.

| Role | Apprentice | VS (x5) | PTA | LTR | QM |
| --- | --- | --- | --- | --- | --- |
| Hourly Rate (£/hr) | 77.80 | 77.80 | 96.20 | 119.30 | 96.20 |
| Time Commitment (hours) | 1200 | 37 | 37 | 14.8 | 14.8 |
| Total cost (£) | 93360 | 14393 | 3559.40 | 1765.64 | 1423.76 |

Table 1 - Time Commitments and Costings for staffing on this project

## Other costs

This project uses a previously purchased instrument in an established lab, as such instrument and lab costs are not required to be covered by this project. This includes all lab consumables such as solvents and analytical columns. The samples and standards used for method development and validation have been previously acquired for different work so again their costs are outside of the scope of this project. The only foreseen acquisitions are Certified Reference Materials (CRMs) of the stabilisers of interest as these will be needed for validation work, the cost for purchasing these is not to exceed £1500 total.

## Time planning

Figure 1 - Gantt Chart of Projected timetable of project

Figure 1 - Gantt Chart of Projected timetable of project

The timetable for the project is to be completed within Financial Year 26/27. The time allocations for the constituent activities are shown in Figure 1.

# References

| [1] | NATO, “AOP-48: EXPLOSIVES, NITROCELLULOSE-BASED PROPELLANTS, STABILITY TEST PROCEDURES AND REQUIREMENTS USING STABILIZER DEPLETION,” NATO, New York, 2008. |
| --- | --- |
| [2] | Eurachem, The Fitness for Purpose of Analytical Methods Third Edition, Eurachem, 2025. |
| [3] | K. Gańczyk-Specjalska, “Conventional and alternative nitrocellulose stabilisers used in gun propellants,” High Energy Materials, vol. 11, no. 2, pp. 73-82, 2019. |
| [4] | W. P. C. de Klerk, “Assessment of Stability of Propellants and Safe Lifetimes,” Propellants, Explosives, Pyrotechnics, vol. 40, no. 3, pp. 388-393, 2015. |
| [5] | D. Trache and A. F. Tarchoun, “Analytical methods for stability assessment of nitrate esters-based propellants,” Critical Reviews in Analytical Chemistry , vol. 49, no. 5, pp. 415-438, 2019. |
| [6] | X. Yao, A. McShane and M. Castillo, “Chapter 17 - Quantitative Proteomics in Development of Disease Protein Biomarkers,” in Proteomic and Metabolomic Approaches to Biomarker Discovery, Academic Press, 2013, pp. 259-278. |
| [7] | Y. P. N. Yang and E. Reeves, “Quantitative Analysis of Organic Gunshot Residues by LC/MS/MS,” Agilent Technologies, 2010. |
| [8] | C. Pyl, W. Feeney, L. Arroyo and T. Trejos, “Capabilities and limitations of GC–MS and LC-MS/MS for trace detection of organic gunshot residues from skin specimens,” Forensic Chemistry, vol. 33, 2023. |
| [9] | S. Bell and W. Feeney, “Single shot, single sample, single instrument detection of IGSR and OGSR using LC/MS/MS,” Forensic Science International, vol. 299, pp. 215-22, 2019. |
| [10] | D. Perret, S. Marchese, A. Gentili and R. Curini, “LC–MS–MS Determination of Stabilizers and Explosives Residues in Hand-Swabs,” Chromatographia, vol. 68, pp. 517-524, 2008. |
| [11] | J. Thomas, D. Lincoln and B. McCord, “Separation and Detection of Smokeless Powder Additives by Ultra Performance Liquid Chromatography with Tandem Mass Spectrometry (UPLC/MS/MS),” Journal of Forensic Sciences, vol. 58, no. 3, 2013. |
| [12] | Z. Wu, Y. Tong, J. Yu and X. Zhang, “Detection of N,N′-diphenyl-N,N′-dimethylurea (methyl centralite) in gunshot residues using MS-MS method,” Analyst, vol. 124, pp. 1563-1567, 1999. |
| [13] | E. Goudsmits, G. Sharples and J. Birkett, “Preliminary classification of characteristic organic gunshot residue compounds,” Science & Justice, vol. 56, no. 6, pp. 421-425, 2016. |
| [14] | D. Laza, B. Nys, J. De Kinder, A. Kirsch-De Mesmaeker and C. Moucheron, “Development of a quantitative LC-MS/MS method for the analysis of common propellant powder stabilizers in gunshot residue,” Journal of Forensic Sciences, vol. 52, no. 4, pp. 842-850, 2007. |

# ANNEX A – UWTSD Ethics Form

APPLICATION FOR ETHICAL APPROVAL

In order for research to result in benefit and minimise risk of harm, it must be conducted ethically.

The University follows the OECD Frascati manual definition of research activity: “creative work undertaken on a systematic basis in order to increase the stock of knowledge, including knowledge of man, culture and society, and the use of this stock of knowledge to devise new applications”. As such this covers activities undertaken by members of staff, postgraduate research students, and both taught postgraduate and undergraduate students working on dissertations/projects.

The individual undertaking the research activity is known as the “principal researcher”.

This form must be completed and approved prior to undertaking any research activity.

SECTION A: About You (Principal Researcher)

| 1 | Full Name: | Quinn Alexander Dunbar |
| --- | --- | --- |
| 2 | Student Number: | 2217666 |
| 3 | Email address: | 2217666@student.uwtsd.ac.uk |
| 4 | Programme of Study: | BEng (Hons) Ordnance, Munitions and Explosives |
| 5 | Director of Studies/Supervisor: | Dr Ewa Kazimierska |

SECTION B:  Internal and External Ethical Guidance Materials

|  | Please list the core ethical guidance documents that have been referred to during the completion of this form (including any discipline-specific codes of research ethics, location-specific dodes of research ethics, and also any specific ethical guidance relating to the proposed methodology). Please tick to confirm that your research proposal adheres to these codes and guidelines. You may add rows to this table if needed. |  |
| --- | --- | --- |
| 1 | UWTSD Research Ethics & Integrity Code of Practice | ☒ |
| 2 | UWTSD Research Data Management Policy | ☒ |
| 3 |  | ☐ |

SECTION C: Details of Research Activity

| 1 | Indicative title: | Development of a Method for Propellant Stabiliser Quantitation using High Performance Liquid Chromatography – Triple Quadrupole Mass Spectrometry |  |  |
| --- | --- | --- | --- | --- |
| 2 | Proposed start date: | June 2026 | Proposed end date: | March 2027 |
|  | Introduction to the Research (maximum 300 words in each section) Ensure that you write for a Non-Specialist Audience when outlining your response to the three points below: - Purpose of Research Activity - Proposed Research Question - Aims of Research Activity - Objectives of Research Activity Demonstrate, briefly, how Existing Research has informed the proposed activity and explain - What the research activity will add to the body of knowledge - How it addresses an area of importance. |  |  |  |
| 3 | Purpose of Research Activity To allow for the identification of quantification of stabiliser compounds in a given sample of a propellant by ultra high pressure liquid chromatography coupled tandem mass spectrometry (UHPLC-MS/MS) |  |  |  |
| 4 | Research Question How can the identification and quantification of propellant stabilisers be improved by the use of UHPLC-MS/MS? |  |  |  |
| 5 | Aims of Research Activity - Develop a method for the identification and quantification of stabilisers by UHPLC-MS/MS - Validate the method to show it is fit for purpose and has good limits of detection and quantification |  |  |  |
| 6 | Objectives of Research Activity - Compose a validated method statement for the identification and quantification of stabilisers by UHPLC-MS/MS |  |  |  |
|  | Proposed data collection methods (maximum 600 words) Provide a brief summary of all the methods that may be used in the research activity to collect data, making it clear what specific techniques may be used. If methods other than those listed in this section are deemed appropriate later, additional ethical approval for those methods will be needed. You do not need to justify the methods here, but should instead describe how you intend to collect the data necessary for you to complete your project. |  |  |  |
| 7 | All data gathered shall be in the form of data produced by the instrument. This will include the participation of other scientists but no data will be gathered directly from them or regarding their behaviour, only data relevant to the analysis of samples will be collected. |  |  |  |

SECTION D: Scope of Research Activity

|  | Will the research activity include: | YES | NO |
| --- | --- | --- | --- |
| 1 | Use of a questionnaire or similar research instrument? | ☐ | ☒ |
| 2 | Use of interviews? | ☐ | ☒ |
| 3 | Use of focus groups? | ☐ | ☒ |
| 4 | Use of participant diaries? | ☐ | ☒ |
| 5 | Use of video or audio recording? | ☐ | ☒ |
| 6 | Use of computer-generated log files? | ☐ | ☒ |
| 7 | Participant observation with their knowledge? | ☐ | ☒ |
| 8 | Participant observation without their knowledge? | ☐ | ☒ |
| 9 | Access to personal or confidential information without the participants’ specific consent? | ☐ | ☒ |
| 10 | Administration of any questions, test stimuli, presentation that may be experienced as physically, mentally or emotionally harmful / offensive? | ☐ | ☒ |
| 11 | Performance of any acts which may cause embarrassment or affect self-esteem? | ☐ | ☒ |
| 12 | Investigation of participants involved in illegal activities? | ☐ | ☒ |
| 13 | Use of procedures that involve deception? | ☐ | ☒ |
| 14 | Administration of any substance, agent or placebo? | ☐ | ☒ |
| 15 | Working with live vertebrate animals? | ☐ | ☒ |
| 16 | Procedures that may have a negative impact on the environment? | ☐ | ☒ |
| 17 | Other primary data collection methods. Please indicate the type of data collection method(s) below. | ☒ | ☐ |
|  | Details of any other primary data collection method: Data to be analysed will be in the form of reports and figures/ values produced by the instrument. In the case of validation experiments involving multiple users the users’ information will be anonymised. |  |  |

If you have ticked NO to every question then the research activity is (ethically) low risk and you may skip section E and continue to section F.

If YES to any question, then no research activity should be undertaken until full ethical approval has been obtained.

SECTION E: Intended Participants

|  | Who are the intended participants: | YES | NO |
| --- | --- | --- | --- |
| 1 | Students or staff at the University? | ☐ | ☒ |
| 2 | Adults (over the age of 18 and competent to give consent)? | ☒ | ☐ |
| 3 | Vulnerable adults? | ☐ | ☒ |
| 4 | Children and Young People under the age of 18? (Consent from Parent, Carer or Guardian will be required) | ☐ | ☒ |
| 5 | Prisoners? | ☐ | ☒ |
| 6 | Young offenders? | ☐ | ☒ |
| 7 | Those who could be considered to have a particularly dependent relationship with the investigator or a gatekeeper? | ☐ | ☒ |
| 8 | People engaged in illegal activities? | ☐ | ☒ |
| 9 | Others. Please indicate the participants below, and specifically any group who may be unable to give consent. | ☒ | ☐ |
|  | Details of any other participant groups: Participants will only be scientists that are trained in the use of the instrumentation. |  |  |

|  | Participant numbers and source Provide an estimate of the expected number of participants. How will you identify participants and how will they be recruited? |  |
| --- | --- | --- |
| 10 | How many participants are expected? | 5 |
| 11 | Who will the participants be? | Other scientists employed by my employer |
| 12 | How will you identify the participants? | Validation scientists will be identified by a number attached to their produced data only. |

|  | Information for participants: | YES | NO | N/A |
| --- | --- | --- | --- | --- |
| 13 | Will you describe the main research procedures to participants in advance, so that they are informed about what to expect? | ☒ | ☐ | ☐ |
| 14 | Will you tell participants that their participation is voluntary? | ☐ | ☒ | ☐ |
| 15 | Will you obtain written consent for participation? | ☐ | ☐ | ☒ |
| 16 | Will you explain to participants that refusal to participate in the research will not affect their treatment or education (if relevant)? | ☐ | ☐ | ☒ |
| 17 | If the research is observational, will you ask participants for their consent to being observed? | ☐ | ☐ | ☒ |
| 18 | Will you tell participants that they may withdraw from the research at any time and for any reason? | ☐ | ☐ | ☒ |
| 19 | With questionnaires, will you give participants the option of omitting questions they do not want to answer? | ☐ | ☐ | ☒ |
| 20 | Will you tell participants that their data will be treated with full confidentiality and that, if published, it will not be identifiable as theirs? | ☐ | ☐ | ☒ |
| 21 | Will you debrief participants at the end of their participation, in a way appropriate to the type of research undertaken? | ☐ | ☐ | ☒ |
| 22 | If NO to any of above questions, please give an explanation |  |  |  |
|  | Validation scientists will assist in this project as part of their normal job role |  |  |  |

|  | Information for participants: | YES | NO | N/A |
| --- | --- | --- | --- | --- |
| 24 | Will participants be paid? | ☐ | ☐ | ☒ |
| 25 | Is specialist electrical or other equipment to be used with participants? | ☐ | ☐ | ☒ |
| 26 | Are there any financial or other interests to the investigator or University arising from this study? | ☐ | ☐ | ☒ |
| 27 | Will the research activity involve deliberately misleading participants in any way, or the partial or full concealment of the specific study aims? | ☐ | ☐ | ☒ |
| 28 | If YES to any question, please provide full details |  |  |  |
|  | N/A |  |  |  |

SECTION F: Anticipated Risks

|  | Outline any anticipated risks that may adversely affect any of the participants, the researchers and/or the University, and the steps that will be taken to address them. |  |
| --- | --- | --- |
| 1 | Risks to participants For example: sector-specific health & safety, emotional distress, financial disclosure, physical harm, transfer of personal data, sensitive organisational information. If you have identified in section D that there are no participants then enter N/A and go skip to question 3. |  |
|  | Risk to participants: Participants will face the usual risks associated with laboratory and instrument use, such as chemicals and moving parts. | How you will mitigate the risk to participants: Following the employer risk assessments covering the use of laboratories. |
| 2 | If research activity may include sensitive, embarrassing or upsetting topics (e.g. sexual activity, drug use) or issues likely to disclose information requiring further action (e.g. criminal activity), give details of the procedures to deal with these issues, including any support/advice (e.g. helpline numbers) to be offered to participants. Note that where applicable, consent procedures should make it clear that if something potentially or actually illegal is discovered in the course of a project, it may need to be disclosed to the proper authorities |  |
|  | N/A |  |
| 3 | Risks to the investigator For example: personal health & safety, physical harm, emotional distress, risk of accusation of harm/impropriety, conflict of interest |  |
|  | Risk to the investigator: The investigator will face the usual risks associated with laboratory and instrument use, such as chemicals and moving parts. | How you will mitigate the risk to the investigator: Following the employer risk assessments covering the use of laboratories. |
| 4 | University/institutional risks For example: adverse publicity, financial loss, data protection |  |
|  | Risk to the University: Adverse publicity may be received by the university for its association with the DSTL and the wider defence sector by groups that target defence. | How you will mitigate the risk to the University: All research and activities within the project will be carried out to Dstl procedures. Any documents for university use will be reviewed as acceptable to release beforehand. The University is to secure any given information in the agreed fashion. |
| 5 | Environmental risks For example: accidental spillage of pollutants, damage to local ecosystems |  |
|  | Risk to the environment: There is the possibility for environmental contamination from the chemicals used in this project. | How you will mitigate the risk to environment: All waste shall be disposed of by the approved waste management pathways. |

SECTION G: Feedback, Consent and Confidentiality

If you have identified in section D that there are no participants then enter skip this section and continue to section H.

| 1 | Feedback What de-briefing and feedback will be provided to participants, how will this be done and when? |
| --- | --- |
|  | Debriefing is not required as it is not the participant’s behaviour that is being reviewed. |
| 2 | Informed consent Describe the arrangements to inform potential participants, before providing consent, of what is involved in participating. Describe the arrangements for participants to provide full consent before data collection begins. If gaining consent in this way is inappropriate, explain how consent will be obtained and recorded in accordance with prevailing data protection legislation. |
|  | Validation scientist will be conducting any work as part of their normal job role. |
| 3 | Confidentiality / Anonymity Set out how anonymity of participants and confidentiality will be ensured in any outputs. If anonymity is not being offered, explain why this is the case. |
|  | Participant identity will be limited to a validation scientist number to denote which datasets were gathered by not the lead investigator |

SECTION H: Data Protection and Storage

|  | Does the research activity involve personal data (as defined by the General Data Protection Regulation 2016 “GDPR” and the Data Protection Act 2018 “DPA”)? | YES | NO |
| --- | --- | --- | --- |
| 1 | “Personal data” means any information relating to an identified or identifiable natural person (‘data subject’). An identifiable natural person is one who can be identified, directly or indirectly, in particular by reference to an identifier such as a name, an identification number, location data, an online identifier or to one or more factors specific to the physical, physiological, genetic, mental, economic, cultural or social identity of that natural person. Any video or audio recordings of participants is considered to be personal data. | ☐ | ☒ |
|  | If YES, provide a description of the data and explain why this data needs to be collected: |  |  |
| 2 |  |  |  |
|  | Does it involve special category data (as defined by the GDPR)? | YES | NO |
| 3 | “Special category data” means sensitive personal data consisting of information as to the data subjects’ – (a) racial or ethnic origin, (b) political opinions, (c ) religious beliefs or other beliefs of a similar nature, (d) membership of a trade union (within the meaning of the Trade Union and Labour Relations (Consolidation) Act 1992), (e) physical or mental health or condition, (f) sexual life, (g) genetics, (h) biometric data (as used for ID purposes), | ☐ | ☒ |
|  | If YES, provide a description of the special category data and explain why this data needs to be collected: |  |  |
| 4 |  |  |  |

|  | Will data from the research activity (collected data, drafts of the thesis, or materials for publication) be stored in any of the following ways? | YES | NO |
| --- | --- | --- | --- |
| 5 | Manual files (i.e. in paper form)? | ☐ | ☒ |
| 6 | University computers? | ☐ | ☒ |
| 7 | Private company computers? | ☒ | ☐ |
| 8 | Home or other personal computers? | ☐ | ☒ |
| 9 | Laptop computers/ CDs/ Portable disk-drives/ memory sticks? | ☒ | ☐ |
| 10 | “Cloud” storage or websites? | ☐ | ☒ |
| 11 | Other – specify: | ☐ | ☒ |
| 12 | For all stored data, explain the measures in place to ensure the security of the data collected, data confidentiality, including details of backup procedures, password protection, encryption, anonymisation and pseudonymisation: |  |  |
|  | Raw data shall be stored and retrieved from the computer controlling the instrument, this is password controlled and not connected to any networks. All written-up data and drafts/ reports will be stored on a secure DSTL network until the final report is approved for release to the university system. |  |  |

|  | Data Protection |  |  |
| --- | --- | --- | --- |
|  | Will the research activity involve any of the following activities: | YES | NO |
| 13 | Electronic transfer of data in any form? | ☒ | ☐ |
| 14 | Sharing of data with others at the University outside of the immediate research team? | ☐ | ☒ |
| 15 | Sharing of data with other organisations? | ☐ | ☒ |
| 16 | Export of data outside the UK or importing of data from outside the UK? | ☐ | ☒ |
| 17 | Use of personal addresses, postcodes, faxes, emails or telephone numbers? | ☐ | ☒ |
| 18 | Publication of data that might allow identification of individuals? | ☐ | ☒ |
| 19 | If YES to any question, please provide full details, explaining how this will be conducted in accordance with the GDPR and Data Protection Act (2018) (and/or any international equivalent): |  |  |
|  | N/A |  |  |
| 20 | List all who will have access to the data generated by the research activity: |  |  |
|  | DSTL (All Data) UWTSD (Final Report) |  |  |
| 21 | List who will have control of, and act as custodian(s) for, data generated by the research activity: |  |  |
|  | Quinn Dunbar (Lead investigator) |  |  |
| 22 | Give details of data storage arrangements, including security measures in place to protect the data, where data will be stored, how long for, and in what form. |  |  |
|  | Data will be stored on secure Dstl computers and network as per Dstl data handling procedures. |  |  |
| 22 | Confirm that you have read the UWTSD guidance on data management (see https://www.uwtsd.ac.uk/library/research-data-management/) | ☒ |  |
| 23 | Confirm that you are aware that you need to keep all data until after your research has completed or the end of your funding | ☒ |  |

SECTION I: Declaration

|  | The information which I have provided is correct and complete to the best of my knowledge. I have attempted to identify any risks and issues related to the research activity and acknowledge my obligations and the rights of the participants. In submitting this application I hereby confirm that I undertake to ensure that the above named research activity will meet the University’s Research Ethics and Integrity Code of Practice which is published on the website: https://www.uwtsd.ac.uk/research/research-ethics/ |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 1 | Signature of applicant: | Q Dunbar | Date: 15/6/26 |  |  |
| 2 |  | Director of Studies/Supervisor: | Dr Ewa Kazimierska |  | Date: |
| 3 |  | Signature: |  |  |  |

FOR INTERNAL USE ONLY:

|  | Ethical approval given |  |  |
| --- | --- | --- | --- |
| 1 | Signature of assessor: |  | Date: |
| 2 | Name: |  |  |
| 3 | Role: |  |  |